# System Subsystem
## System Time (:SYSTem:TIME)
Command / Format	:SYSTem:TIME <time> / :SYSTem:TIME?
Instruction	Sets System time. Gets System time.
Parameter Type	String
Parameter Range	hour(0~23), minute(0~59), second(0~59)
Return	String
Default	None
Menu	System > date & time
Example	Sets System time: / :SYSTem:TIME 182559 / Gets System time: / :SYSTem:TIME?

## System Date (:SYSTem:DATE)
Command / Format	:SYSTem:DATE <date> / :SYSTem:DATE?
Instruction	Sets system date. Gets system date.
Parameter Type	String
Parameter Range	year(four digits), month(1~12), date(1~31)
Return	String
Default	None
Menu	System > date&time
Example	Sets System date: / :SYSTem:DATE 20050101 / Gets System date: / :SYSTem:DATE?

## IP Address (:SYSTem:COMMunicate:LAN:IPADdress)
Command / Format	:SYSTem:COMMunicate:LAN:IPADdress <address> / :SYSTem:COMMunicate:LAN:IPADdress?
Instruction	Sets a host name for the analyzer in network. Gets IP address.
Parameter Type	String
Parameter Range	Conform to the IP Sets standard(0-255:0-255:0-255:0-255)
Return	IP address String
Default	None
Menu	System > Interface > LAN > IP Address
Example	:SYSTem:COMMunicate:LAN:IPADdress “192.168.1.12” / :SYSTem:COMMunicate:LAN:IPADdress?

## Gateway (:SYSTem:COMMunicate:LAN:GATeway)
Command / Format	:SYSTem:COMMunicate:LAN:GATeway <gateway> / :SYSTem:COMMunicate:LAN:GATeway?
Instruction	Sets the gateway for the analyzer in the network. The gateway will be fetched automatically if the IP assignment is set to DHCP. Gets gateway.
Parameter Type	String
Parameter Range	Conform to the IP standard (0-255:0-255:0-255:0-255)
Return	Gateway String.
Default	None
Menu	System > Interface > LAN > Gateway
Example	:SYSTem:COMMunicate:LAN:GATeway “192.168.1.1” / :SYSTem:COMMunicate:LAN:GATeway?

## Subnet Mask (:SYSTem:COMMunicate:LAN:SMASk)
Command / Format	:SYSTem:COMMunicate:LAN:SMASk <mask> / :SYSTem:COMMunicate:LAN:SMASk?
Instruction	Sets the subnet mask according to the PC network Settings. The subnet mask will be set automatically if the IP assignment is set to DHCP.
Parameter Type	String
Parameter Range	Conform to the IP standard (0-255:0-255:0-255:0-255)
Return	Subnet mask string.
Default	None
Menu	System > Interface > LAN > Subnet Mask
Example	:SYSTem:COMMunicate:LAN:SMASk?

## IP Config (:SYSTem:COMMunicate:LAN:TYPE)
Command / Format	:SYSTem:COMMunicate:LAN:TYPE STATIC|DHCP / :SYSTem:COMMunicate:LAN:TYPE?
Instruction	Toggles the IP assignment Setting between static (manual) and DHCP / (dynamic assignment) mode. / Gets IP config.
Parameter Type	Enumeration
Parameter Range	STATIC|DHCP
Return	Enumeration
Default	None
Menu	System > Interface > LAN > IP Config
Example	:SYSTem:COMMunicate:LAN:TYPE DHCP / :SYSTem:COMMunicate:LAN:TYPE?

## Language (:SYSTem:LANGuage)
Command / Format	:SYSTem:LANGuage SCHINESE|ENGLISH / :SYSTem:LANGuage?
Instruction	Sets language. / Gets language.
Parameter Type	Enumeration
Parameter Range	SCHINESE: Chinese / ENGLISH: English
Return	Enumeration
Default	None
Menu	System > Language
Example	Sets language / :SYSTem:LANGuage SCHINESE / Gets language / :SYSTem:LANGuage?

## Power On Type (:SYSTem:PON:TYPE)
Command / Format	:SYSTem:PON:TYPE DFT|LAST|USER / :SYSTem:PON:TYPE?
Instruction	Uses command to set analyzer to power on in default, user, or last state. / Gets power on type.
Parameter Type	Enumeration
Parameter Range	DFT: Default / LAST: Last / USER: Custom Configuration
Return	Enumeration
Default	DFT
Menu	System > Pwr/Preset > Power On
Example	SYSTem:PON:TYPE DFT

## System Preset (:SYSTem:PRESet)
Command / Format	:SYSTem:PRESet
Instruction	Use this command to preset the instrument. The preset type is based on the Setting of Preset Type: DFT, User or Last.
Parameter Type	None
Parameter Range	None
Return	None
Default	None
Menu	None
Example	:SYSTem:PRESet

## System Restart (:SYSTem:RESTart)
Command / Format	:SYSTem:RESTart
Instruction	Use this command to restart the instrument (part of machine may not support).
Parameter Type	None
Parameter Range	None
Return	None
Default	None
Menu	None
Example	:SYSTem:RESTart

## Preset Type (:SYSTem:PRESet:TYPE)
Command / Format	:SYSTem:PRESet:TYPE DFT|LAST|USER / :SYSTem:PRESet:TYPE?
Instruction	Uses this command to preset the analyzer to default, user, or last state. / Gets preset type.
Parameter Type	Enumeration
Parameter Range	DFT: Default / LAST: Last / USER: Custom Configuration
Return	Enumeration
Default	DFT
Menu	System > Pwr/Preset > Preset
Example	:SYSTem:PRESet:TYPE DFT

## Factory ReSet (:SYSTem:FDEFault)
Command / Format	:SYSTem:FDEFault
Instruction	Sets both the measure and setting parameters to factory preset parameters
Parameter Type	None
Parameter Range	None
Return	None
Default	None
Menu	System > Pwr/Preset > Factory Reset
Example	:SYSTem:FDEFault

## Enable Option (:SYSTem:LKEY)
Command / Format	:SYSTem:LKEY <option>,<license key>
Instruction	Use this command to enable the specified option with the license key, please restart the instrument to make license active.
Parameter Type	“option”: Enumeration / “ license key”: String
Parameter Range	“option”: Meas|EMI|CAT|TG / “ license key”: provided by Siglent Technologies, 16 bits String.
Return	None
Default	None
Menu	System > System Info > Load Option
Example	:SYSTem:LKEY EMI,fjbdajffnklmgwno

## Installed Options Query (:SYSTem:OPTions?)
Command / Format	:SYSTem:OPTions?
Instruction	This command returns a list of the options that are installed.
Parameter Type	None
Parameter Range	None
Return	Meas| EMI|CAT|TG
Default	None
Menu	System > System Info
Example	:SYSTem:OPTions?

## Power Off (:SYSTem:POWer:OFF)
Command / Format	:SYSTem:POWer:OFF
Instruction	Use this command to turn off the instrument.
Parameter Type	None
Parameter Range	None
Return	None
Default	None
Menu	None
Example	:SYSTem:POWer:OFF

## Screenshot Data (:HCOPy:SDUMp:DATA?)
Command / Format	:HCOPy:SDUMp:DATA?
Instruction	Use this command to query the screenshot data(.bmp)
Parameter Type	None
Parameter Range	None
Return	String
Default	None
Menu	System > System Info
Example	:HCOPy:SDUMp:DATA?

## System Info (:SYSTem:CONFigure:SYSTem?)
Command / Format	:SYSTem:CONFigure:SYSTem?
Instruction	Use this command to query the system message of the instrument
Parameter Type	None
Parameter Range	None
Return	String
Default	None
Menu	System > System Info
Example	:SYSTem:CONFigure:SYSTem?