# SYSTem Commands

SYSTem Commands are used for the basic operations of an oscilloscope: RUN/STOP control, operation of the error queue and system setup data. SYSTem Commands include:
- :RUN
- :STOP
- :AUTO
- :SYSTem:ERRor
- :SYSTem:SETup
We will give detailed introductions for each command in the following parts.
## :RUN
Command format	:RUN
Function	Execute this command, the oscilloscope will start waveform sampling working. / To stop working, execute :STOP command again.

## :STOP
Command format	:STOP
Function	Execute this command, the oscilloscope will stop waveform sampling working. / To restart working, execute :RUN command again.

## :AUTO
Command format	:AUTO
Function	The command makes the oscilloscope tests all input waveforms and set the waveforms automatically to get the optimum conditions to display.

## :SYSTem:ERRor
Command format	:SYSTem:ERRor / :SYSTem:ERRor?
Function	The command clears the queue of error information.
Returned format	The query returns the last error, such as “Undefined header”. If there is no error, return “0, No error”.

System Error Code
Up to 10 errors can be recorded in the system error queue. If not enough, the system will adopt FIFO manner to cover the original error record.

The SYST:ERR? Command is used to read the first error code in the form of “error code, error description”, so as to reduce the error number of error queue. For instance, if no error appears, the system will return: 0, No error.

Besides, the :SYST:ERR Command is able to clear error queue.

Error Code	Mnemonic Symbol	Error Description
0	ERR_NONE	No error
1	ERR_SAME_SETTING	Same setting
2	ERR_INVALID_INPUT	Invalid input
3	ERR_LIMIT_SETTING	Setting limit
4	ERR_CH_OFFSET_LIMIT	Channel offset limit
5	ERR_CH_SCALE_LIMIT	Channel scale limit
6	ERR_CH_PROBE_LIMIT	Channel probe limit
7	ERR_CH_FILTER_LIMIT	Channel filter limit
8	ERR_TIME_OFFSET_LIMIT	Timebase offset limit
9	ERR_TIME_SCALE_LIMIT	Timebase scale limit
10	ERR_TIME_DELAYED_OFFSET_LIMIT	Timebase of timedelay offset limit
11	ERR_TIME_DELAYED_SCALE_LIMIT	Timebase of timedelay scale limit
12	ERR_TRIG_LEVEL_LIMIT	Trigger level limit
13	ERR_MATH_VERT_OFFSET_LIMIT	Math vertical offset limit
14	ERR_MATH_VERT_SCALE_LIMIT	Math vertical scale limit
15	ERR_FFT_VERT_SCALE_LIMIT	FFT vertical offset limit
16	ERR_FFT_VERT_OFFSET_LIMIT	FFT vertical offset limit
17	ERR_FFT_HORIZ_SCALE_LIMIT	FFT horizontal scale limit
18	ERR_FFT_HORIZ_OFFSET_LIMIT	FFT horizontal offset limit
19	ERR_CUR_A_X_LIMIT	CursorA X-Axial limit
20	ERR_CUR_B_X_LIMIT	CursorB X-Axial limit
21	ERR_CUR_A_Y_LIMIT	CursorA Y-Axial limit
22	ERR_CUR_B_Y_LIMIT	CursorB Y-Axial limit
23	ERR_HOLDOFF_TIME_LIMIT	Holdoff time limit
24	ERR_INTENSITY_LIMIT	Intensity limit
25	ERR_PULSE_WIDTH_LIMIT	Pulse width limit
26	ERR_VIDEO_LINE_LIMIT	Video line limit
27	ERR_REC_INTERVAL_LIMIT	Record interval limit
28	ERR_REC_END_FRAME_LIMIT	Record end frame limit
29	ERR_PLAY_INTERVAL_LIMIT	Play interval limit
30	ERR_PLAY_START_FRAME_LIMIT	Play start frame limit
31	ERR_PLAY_CUR_FRAME_LIMIT	Play current frame limit
32	ERR_PLAY_END_FRAME_LIMIT	Play end frame limit
33	ERR_STOARAGE_START_FRAME_LIMIT	Storage start frame limit
34	ERR_STOARGE_END_FRAME_LIMIT	Storage end frame limit
35	ERR_REF_VERT_OFFSET_LIMIT	Ref vertical offset limit
36	ERR_REF_VERT_SCALE_LIMIT	Ref vertical scale limit
37	ERR_PF_MASK_LIMIT	Passfail mask limit
38	ERR_SAMPLING_RATE_LIMIT	Sampling rate limit
39	ERR_GRID_INTENSITY_LIMIT	Grid intensity limit
40	ERR_TRIG_SENSITIVITY_LIMIT	Trigger sensitivity limit
41	ERR_TRIG_SLOPE_TIME_LIMIT	Trigger slop time limit
42	ERR_MEM_DEPTH_LIMIT	Memory depth limit
43	ERR_FUNCTION_NOT_AVAILABLE	Function not available
44	ERR_LOCATION_EMPTY	Location empty
45	ERR_MEAS_ALREADY_SELECTED	Measure already selected
46	ERR_NO_SIGNAL_FOUND	No signal found
47	ERR_WAVEFORM_RECORD_FINISHED	Waveform record finished
48	ERR_FILE_UTILITY_FAIL	File utility fail
49	ERR_CHANNEL_INVALID	Channel invalid
50	ERR_AUTO_KEY_LIMITED	Auto key limited
51	ERR_NOT_ENOUGH_MEMORY	Not enough memory
52	ERR_WAVE_SAVE_FAILED	Waveform save failed
53	ERR_WAVE_LOAD_FAILED	Waveform load failed
54	ERR_FILE_IS_COVERED	File is covered
55	ERR_FILTER_IS_CLOSED	Filter is closed
56	ERR_WAVE_TYPE_NONE	No signal detected
57	ERR_WAVE_TYPE_DC	DC signal detected
58	ERR_WAVE_TYPE_SINE	Sine signal detected
59	ERR_WAVE_TYPE_RAMP	Triangle signal detected
60	ERR_WAVE_TYPE_RECT	Square signal detected
61	ERR_WAVE_TYPE_UNKNOWN	Unknown signal detected
62	CMD_ERR	Error header
63	CMD_NOT_PARSE	Undefined header
64	ERR_PF_OUTPUT	PassFail Out
65	ERR_MISSING_HW	Missing Hardware
66	ERR_OUT_OF_RANGE Out of range
67	ERR_CANNOT_EXECURE Can’t execute

## :SYSTem:SETup
Command format	:SYSTem:SETup <setup_data> / :SYSTem:SETup?
Function	The command downloads the system setup data. <setup data> is a binary data that meets IEEE 488.2 # format.
Returned format	The query returns the value of system setup data.