# Screen commands
## :AUToscale
Syntax	:AUToscale
Description	Enable the waveform auto setting function. The oscilloscope will automatically adjust the vertical scale, horizontal timebase and trigger mode according to the input signal to realize optimum waveform display. This command is equivalent to pressing the AUTO key at the front panel.
Explanation	Theoretically, waveform auto setting function requires that the frequency of sine is no lower than 41Hz; the duty cycle should be greater than 1% and the amplitude must be at least 20mVpp for square (the probe ratio is 1X). / When the pass/fail function is enabled (see the :MASK:ENABle command), if you sent this command, the oscilloscope will disable the pass/fail function firstly and then execute the waveform auto setting function. / When the waveform record function is enabled or during the playback of the recorded waveform, this command is invalid.

## :CLEar
Syntax	:CLEar
Description	Clear all the waveforms on the screen. If the oscilloscope is in the RUN state, waveform will still be displayed. This command is equivalent to pressing the CLEAR key at the front panel.
Related command	:DISPlay:CLEar

## :RUN
## :STOP
Syntax	:RUN / :STOP
Description	The :RUN command starts the oscilloscope and the :STOP command stops the oscilloscope. These commands are equivalent to pressing the RUN/STOP key at the front panel.
Explanation	When the waveform record function is enabled or during the playback of the recorded waveform, these commands are invalid.

## :SINGle
Syntax	:SINGle
Description	Set the oscilloscope to the single trigger mode. This command is equivalent to any of the following two operations: pressing the SINGLE key at the front panel and sending the :TRIGger:SWEep SINGle command.
Explanation	In the single trigger mode, the oscilloscope triggers once when the trigger conditions are met and then stops. / When the waveform record function is enabled or during the playback of the recorded waveform, this command is invalid.
Related commands	:TFORce / :RUN / :STOP

## :TFORce
Syntax	:TFORce
Description	Generate a trigger signal forcefully. This command is only applicable to the normal and single trigger modes (see the :TRIGger:SWEep command) and is equivalent to pressing the FORCE key in the trigger control area at the front panel.