# :CHANnel<n> Commands
The :CHANnel<n> commands are used to set or query the vertical system parameters of the analog channels, such as the bandwidth limit, coupling, vertical scale and vertical offset.

Command List:
- :CHANnel<n>:BWLimit
- :CHANnel<n>:COUPling
- :CHANnel<n>:DISPlay
- :CHANnel<n>:INVert
- :CHANnel<n>:OFFSet
- :CHANnel<n>:RANGe
- :CHANnel<n>:TCAL
- :CHANnel<n>:SCALe
- :CHANnel<n>:PROBe
- :CHANnel<n>:UNITs
- :CHANnel<n>:VERNier
## :CHANnel<n>:BWLimit
Syntax	:CHANnel<n>:BWLimit <type> / :CHANnel<n>:BWLimit?
Description	Set or query the bandwidth limit parameter of the specified channel.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<type>	Discrete	{20M|OFF}	OFF
Explanation	OFF: disable the bandwidth limit and the high frequency components of the signal under test can pass the channel. / 20M: enable the bandwidth limit and the high frequency components of the signal under test that exceed 20 MHz are attenuated. / Enabling the bandwidth limit can reduce the noise, but can also attenuate the high frequency components.
Return format	The query returns 20M or OFF.
Example	:CHANnel1:BWLimit 20M /*Enable the 20MHz bandwidth limit*/ / :CHANnel1:BWLimit? /*The query returns 20M*/

## :CHANnel<n>:COUPling
Syntax	:CHANnel<n>:COUPling <coupling> / :CHANnel<n>:COUPling?
Description	Set or query the coupling mode of the specified channel.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<coupling>	Discrete	{AC|DC|GND}	DC
Explanation	AC: the DC components of the signal under test are blocked. / DC: the DC and AC components of the signal under test can both pass the channel. / GND: the DC and AC components of the signal under test are both blocked.
Return format	The query returns AC, DC or GND.
Example	:CHANnel1:COUPling AC /*Select the AC coupling mode*/ / :CHANnel1:COUPling? /*The query returns AC*/

## :CHANnel<n>:DISPlay
Syntax	:CHANnel<n>:DISPlay <bool> / :CHANnel<n>:DISPlay?
Description	Enable or disable the specified channel or query the status of the specified channel.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<bool>	Bool	{{1|ON}|{0|OFF}}	CH1: 1|ON / CH2 to CH4: 0|OFF
Explanation	The query returns 1 or 0.
Return format	:CHANnel1:DISPlay ON /*Enable CH1*/ / :CHANnel1:DISPlay? /*The query returns 1*/
Example

## :CHANnel<n>:INVert
Syntax	:CHANnel<n>:INVert <bool> / :CHANnel<n>:INVert?
Description	Enable or disable the inverted display mode of the specified channel or query the status of the inverted display mode of the specified channel.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Explanation	When the inverted display mode is enabled, the oscilloscope displays the waveform inverted taking the vertical offset horizontal line as reference.
Return format	The query returns 1 or 0.
Example	:CHANnel1:INVert ON /*Enable the inverted display mode of CH1*/ / :CHANnel1:INVert? /*The query returns 1*/

## :CHANnel<n>:OFFSet
Syntax	:CHANnel<n>:OFFSet <offset> / :CHANnel<n>:OFFSet?
Description	Set or query the vertical offset of the specified channel. The default unit is V.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<offset>	Real	Related to the current vertical scale and probe ratio / When the probe ratio is 1X, / vertical scale≥500mV/div: -100V to +100V vertical scale<500mV/div: -2V to +2V When the probe ratio is 10X, / vertical scale≥5V/div: -1000V to +1000V vertical scale<5V/div: -20V to +20V	0V (the probe ratio is 10X)
Explanation	The query returns the vertical offset in scientific notation.
Return format	:CHANnel1:OFFSet 0.01 /*Set the vertical offset of CH1 to 10mV*/ / :CHANnel1:OFFSet? /*The query returns 1.000000e-02*/
Example

## :CHANnel<n>:PROBe
Syntax	:CHANnel<n>:PROBe <atten> / :CHANnel<n>:PROBe?
Description	Set or query the probe ratio of the specified channel.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<atten>	Discrete	{0.01|0.02|0.05|0.1|0.2|0.5|1|2|5|10|20|50| / 100|200|500|1000}	10
Explanation	Setting the probe ratio refers to multiply the signal sampled with the specified ratio and then display the result (the actual amplitude of the signal will not be affected). / Setting the probe ratio will affect the range of the vertical scale.
Return format	The query returns the probe ratio in scientific notation.
Example	:CHANnel1:PROBe 10 /*Set the probe ratio of CH1 to 10X*/ / :CHANnel1:PROBe? /*The query returns 1.000000e+01*/
Related command	:CHANnel<n>:SCALe

## :CHANnel<n>:RANGe
Syntax	:CHANnel<n>:RANGe <range> / :CHANnel<n>:RANGe?
Description	Set or query the vertical range of the specified channel. The default unit is V.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<range>	Real	Related to the probe ratio / When the probe ratio is 1X: 8mV to 80V / When the probe ratio is 10X: 80mV to 800V	8V (the probe ratio is 10X)
Explanation	This command indirectly modifies the vertical scale of the specified channel (vertical scale = vertical range/8). The vertical scale can be set by the :CHANnel<n>:SCALe command.
Return format	The query returns the vertical range in scientific notation.
Example	:CHANnel1:RANGe 8 /*Set the vertical range of CH1 to 8V*/ / :CHANnel1:RANGe? /*The query returns 8.000000e+00*/

## :CHANnel<n>:SCALe
Syntax	:CHANnel<n>:SCALe <scale> / :CHANnel<n>:SCALe?
Description	Set or query the vertical scale of the specified channel. The default unit is V.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<scale>	Real	Related to the current probe ratio / When the probe ratio is 1X: 1mV to 10V / When the probe ratio is 10X (default): 10mV to / 100V	1V (the probe ratio is 10X)
Explanation	range of the vertical scale is related to the current probe ratio (set by the :CHANnel<n>:PROBe command). / You can use the :CHANnel<n>:VERNier command to enable or disable the fine adjustment of the vertical scale. By default, the fine adjustment is off. At this point, you can only set the vertical scale in 1-2-5 step, namely 10mV, 20mV, 50mV, 100mV, …, 100V (the probe ratio is 10X). When the fine adjustment is on, you can further adjust the vertical scale within a relatively smaller range to improve the vertical resolution. If the amplitude of the input waveform is a little bit greater than the full scale under the current scale and the amplitude would be a little bit lower if the next scale is used, fine adjustment can be used to improve the display amplitude of the waveform to view the signal details.
Return format	The query returns the vertical scale in scientific notation.
Example	:CHANnel1:SCALe 1 /*Set the vertical scale of CH1 to 1V*/ / :CHANnel1:SCALe? /*The query returns 1.000000e+00*/

## :CHANnel<n>:TCAL
Syntax	:CHANnel<n>:TCAL <val> / :CHANnel<n>:TCAL?
Description	Set or query the delay calibration time of the specified channel to calibrate the zero offset of the corresponding channel. The default unit is s.
Parameter	Name		Type	Range	Default
	<n>		Discrete	{1|2|3|4}	--
	<val>		Real	-100ns to 100ns	0.00s
Explanation	<val> can only be set to the specific values in the specified step. If the parameter you / sent is not one of the specific values, the parameter will be set to the nearest specific values automatically. The step varies with the horizontal timebase (set by the :TIMebase[:MAIN]:SCALe command), as shown in the table below.
		Horizontal Timebase	Step of the Delay Calibration Time
		5ns	100ps
		10ns	200ps
		20ns	400ps
		50ns	1ns
		100ns	2ns
		200ns	4ns
		500ns	10ns
		1μs to 10μs	20ns
		Note: When the horizontal timebase is equal to or greater than 10μs, the delay calibration time cannot be adjusted.
Return format	The query returns the delay calibration time in scientific notation.
Example	:CHANnel1:TCAL 0.00000002 /*Set the delay calibration time to 20ns*/ / :CHANnel1:TCAL? /*The query returns 2.000000e-08*/

## :CHANnel<n>:UNITs
Syntax	:CHANnel<n>:UNITs <units> / :CHANnel<n>:UNITs?
Description	Set or query the amplitude display unit of the specified channel.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<units>	Discrete	{VOLTage|WATT|AMPere|UNKNown}	VOLTage
Explanation	The query returns VOLT, WATT, AMP or UNKN.
Return format	:CHANnel1:UNITs VOLTage /*Set the amplitude display unit of CH1 to V*/ / :CHANnel1:UNITs? /*The query returns VOLT*/
Example

## :CHANnel<n>:VERNier
Syntax	:CHANnel<n>:VERNier <bool> / :CHANnel<n>:VERNier?
Description	Enable or disable the fine adjustment of the vertical scale of the specified channel, or query the fine adjustment status of the vertical scale of the specified channel.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2|3|4}	--
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Explanation	By default, the fine adjustment is off. At this point, you can only set the vertical scale in / 1-2-5 step, namely 10mV, 20mV, 50mV, 100mV…100V (the probe ratio is 10X). When the fine adjustment is on, you can further adjust the vertical scale within a relatively smaller range to improve the vertical resolution. If the amplitude of the input waveform is a little bit greater than the full scale under the current scale and the amplitude would be a little bit lower if the next scale is used, fine adjustment can be used to improve the display amplitude of the waveform to view the signal details.
Return format	The query returns 1 or 0.
Example	:CHANnel1:VERNier ON /*Enable the fine adjustment function of the vertical scale of CH1*/ / :CHANnel1:VERNier? /*The query returns 1*/