# :LA Commands (Only for MSO1000Z/MSO1000Z-S)

The :LA commands are used to perform the related operations on the digital channels.

Command List:
- :LA:ACTive
- :LA:AUTosort
- :LA:DIGital<n>:DISPlay
- :LA:DIGital<n>:POSition
- :LA:DIGital<n>:LABel
- :LA:DISPlay
- :LA:POD<n>:DISPlay
- :LA:POD<n>:THReshold
- :LA:SIZE
- :LA:STATe
- :LA:TCALibrate
## :LA:ACTive
Syntax	:LA:ACTive {<digital>|<group>|none} / :LA:ACTive?
Description	Set or query the current active channel or channel group.
Parameter	Name	Type	Range	Default
	<digital>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7| / D8|D9|D10|D11|D12|D13|D14|D15}	D0
	<group>	Discrete	{GROUP1|GROUP2|GROUP3|GROUP4}	--
Explanation	<digital> is used to select any of channels D0-D15. The channel label and waveform corresponding to the channel selected are displayed in red. / <group> is used to select any of user-defined channel groups (GROUP1-GROUP4). All the channel labels and waveforms of the channel group selected are displayed in red. / When none is sent, no channel/channel group will be selected. / Only channels or user-defined channel groups currently turned on can be selected. Please refer to the :LA:DIGital<n>:DISPlay or :LA:DISPlay command to turn on the desired channel or channel group.
Return format	The query returns the current active channel (D0, D1,……or D15), the current active channel group (GROUP1, GROUP2, GROUP3 or GROUP4) or NONE.
Example	:LA:ACTive D3 /*Set the current active channel to D3*/ / :LA:ACTive? /*The query returns D3*/

## :LA:AUTosort
Syntax	:LA:AUTosort <n>
Description	Set the auto ordering mode of the waveforms of the channels turned on on the screen.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{0|1}	0
Explanation	n=1: waveforms on the screen are D0-D15 from top to bottom / n=0: waveforms on the screen are D15-D0 from top to bottom

## :LA:DIGital<n>:DISPlay
Syntax	:LA:DIGital<n>:DISPlay <bool> / :LA:DIGital<n>:DISPlay?
Description	Turn on or off the specified digital channel, or query the status of the specified digital channel.
Parameter	Name	Type	Range	Default
	<n>	Integer	0 to 15	--
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Explanation	You can set the channels currently turned on as the active channel using the :LA:ACTive command.
Return format	The query returns 1 or 0.
Example	:LA:DIGital3:DISPlay ON /*Turn D3 on*/ / :LA:DIGital3:DISPlay? /*The query returns 1*/

## :LA:DIGital<n>:POSition
Syntax	:LA:DIGital<n>:POSition <position> / :LA:DIGital<n>:POSition?
Description	Set or query the display position of the specified digital channel waveform on the screen.
Parameter	Name	Type	Range	Default
	<n>	Integer	0 to 15	--
	<position>	Integer	When the waveform display mode is SMALl: 0 to 15 / When the waveform display mode is LARGe: 0 to 7 / For the waveform display mode, refer to the :LA:SIZE command	--
Explanation	<position> (0 to 15 or 0 to 7) denotes the display position of the waveform on the screen. When the waveform display mode is SMALl, 16 waveforms can be displayed on the screen marked as 0 to 15 from top to bottom. When the waveform display mode is LARGe, 8 waveforms can be displayed on the screen marked as 0 to 7 from top to bottom. / The setting command is only valid when the specified digital channel is turned on.
Return format	The query returns an integer from 0 to 15 or from 0 to 7.
Example	:LA:DIGital1:POSition 3 /*Set the display position of D1 to 3*/ / :LA:DIGital1:POSition? /*The query returns 3*/

## :LA:DIGital<n>:LABel
Syntax	:LA:DIGital<n>:LABel <label> / :LA:DIGital<n>:LABel?
Description	Set or query the label of the specified digital channel.
Parameter	Name	Type	Range	Default
	<n>	Integer	0 to 15	--
	<label>	ASCII string	It can include English uppercase letters (A to Z) and numbers (0 to 9). / It cannot exceed 4 characters.	--
Return format	The query returns the label of the specified digital channel in ASCII string. If the specified channel has no label, the query returns “No Label”.
Example	:LA:DIGital0:LABel ACK /*Set the label of D0 to ACK*/ / :LA:DIGital0:LABel? /*The query returns ACK*/

## :LA:DISPlay
Syntax	:LA:DISPlay {<digital>|<group>|<pod>},<bool> / :LA:DISPlay? {<digital>|<group>|<pod>}
Description	Turn on or off the specified digital channel, user-defined channel group or default channel group, or query the status of the specified digital channel, user-defined channel group or default channel group.
Parameter	Name	Type	Range	Default
	<digital>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7| / D8|D9|D10|D11|D12|D13|D14|D15}	--
	<group>	Discrete	{GROUP1|GROUP2|GROUP3|GROUP4}	--
	<pod>	Discrete	{POD1|POD2}	--
	<bool>	Bool	{{1|ON}|{0|OFF}}	--
Explanation	When any of channels D0 to D7 is turned on, the CH4 function is disabled automatically. When channels D0 to D7 are all turned off, the CH4 function is enabled automatically. / When any of channels D8 to D15 is turned on, the CH3 function is disabled automatically. When channels D8 to D15 are all turned off, the CH4 function is enabled automatically. / <group> denotes the user-defined channel groups. Users can set the digital channels included in the specified user-defined channel group. / Note: A channel can only be added to an unique user-defined channel group. / <pod> denotes the default channel groups: POD1 (D0 to D7) and POD2 (D8 to D15). / The channels and user-defined channel groups currently turned on can be selected as the active channel or channel group using the :LA:ACTive command.
Return format	The query returns 1 or 0.
Example	:LA:DISPlay pod1,ON /*Turn POD1 (D0~D7) on*/ / :LA:DISPlay? pod1 /*The query returns 1*/

## :LA:POD<n>:DISPlay
Syntax	:LA:POD<n>:DISPlay <bool> / :LA:POD<n>:DISPlay?
Description	Turn on or off the specified default channel group, or query the status of the specified default channel group.
Parameter	Name	Type	Range	Default
	<n>	Integer	1 to 2	--
	<bool>	Bool	{{1|ON}|{0|OFF}}	--
Explanation	There are 2 default channel groups: POD1 (D0 to D7) and POD2 (D8 to D15). / When POD1 is turned on, the CH4 function is disabled automatically. When POD1 is turned off, the CH4 function is enabled automatically. / When POD2 is turned on, the CH3 function is disabled automatically. When POD2 is turned off, the CH3 function is enabled automatically.
Return format	The query returns 1 or 0.
Example	:LA:POD1:DISPlay 1 /*Turn POD1 (D0~D7) on*/ / :LA:POD1:DISPlay? /*The query returns 1*/

## :LA:POD<n>:THReshold
Syntax	:LA:POD<n>:THReshold <thre> / :LA:POD<n>:THReshold?
Description	Set or query the threshold of the specified channel group. The default unit is V.
Parameter	Name	Type	Range	Default
	<n>	Integer	1 to 2	--
	<thre>	Real	-15.0V to +15.0V	1.40V
Explanation	There are 2 default channel groups: POD1 (D0 to D7) and POD2 (D8 to D15).
Return format	The query returns the current threshold of the specified channel group in scientific notation.
Example	:LA:POD1:THReshold 3.3 /*Set the threshold of POD1 (D0~D7) to 3.3V*/ / :LA:POD1:THReshold? /*The query returns 3.300000e+00*/

## :LA:SIZE
Syntax	:LA:SIZE <size> / :LA:SIZE?
Description	Set or query the display size of the waveforms of the channels turned on on the screen.
Parameter	Name	Type	Range	Default
	<size>	Discrete	{SMALl|LARGe}	SMALl
Explanation	SMALl: up to 16 waveforms can be displayed on the screen (0 to 15 from top to bottom). / LARGe: up to 8 waveforms can be displayed on the screen (0 to 7 from top to bottom). / The LARGe display mode is only available when the number of channels currently turned on is no more than 8.
Return format	The query returns SMAL or LARG.
Example	:LA:SIZE SMALl /*Set the waveform display size to small*/ / :LA:SIZE? /*The query returns SMAL*/
Related command	:LA:DIGital<n>:POSition

## :LA:STATe
Syntax	:LA:STATe <bool> / :LA:STATe?
Description	Turn on or off the LA function, or query the status of the LA function.
Parameter	Name	Type	Range	Default
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Return format	The query returns 1 or 0.
Example	:LA:STATe ON /*Turn on the LA function*/ / :LA:STATe? /*The query returns 1*/

## :LA:TCALibrate
Syntax	:LA:TCALibrate <tcal> / :LA:TCALibrate?
Description	Set or query the delay calibration time of the digital channels. The default unit is s.
Parameter	Name		Type	Range	Default
	<tcal>		Real	-100ns to 100ns	0.00s
Explanation	When using the oscilloscope for actual measurement, the transmission delay of the probe cable may cause relatively larger error (zero offset). Zero offset is defined as the offset of the crossing point of the waveform and threshold level line relative to the trigger position. For MSO1000Z, users can set a delay time to calibrate the zero offset of the corresponding channel. / You can only set <tcal> to a specified value at certain step. If the parameter sent is not one of the specified values, the specified value nearest to the value sent will be used. The step of this parameter differs in different horizontal timebase (set by the :TIMebase[:MAIN]:SCALe command), as shown in the table below.
		Horizontal Timebase		Delay Calibration Time Step
		5ns		100ps
		10ns		200ps
		20ns		400ps
		50ns		1ns
		100ns		2ns
		200ns		4ns
		500ns		10ns
		1μs to 10μs		20ns
		Note: When the horizontal timebase is equal to or greater than 10μs, the delay calibration time cannot be adjusted.
Return format	The query returns the delay calibration time in scientifc notation.
Example	:LA:TCALibrate 0.00000002 /*Set the delay calibration time to 20ns*/ / :LA:TCALibrate? /*The query returns 2.000000e-08*/