# :MASK Commands

The :MASK commands are used to set and query the pass/fail test parameters.

Command List:
- :MASK:ENABle
- :MASK:SOURce
- :MASK:OPERate
- :MASK:MDISplay
- :MASK:SOOutput
- :MASK:OUTPut
- :MASK:X
- :MASK:Y
- :MASK:CREate
- :MASK:PASSed?
- :MASK:FAILed?
- :MASK:TOTal?
- :MASK:RESet
## :MASK:ENABle
Syntax	:MASK:ENABle <bool> / :MASK:ENABle?
Description	Enable or disable the pass/fail test or query the status of the past/fail test.
Parameter	Name	Type	Range	Default
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Explanation	The pass/fail test is invalid in the following conditions: the horizontal timebase is XY or ROLL, in the slow sweep mode (namely when the horizontal timebase is YT, the horizontal timebase is set to 200ms/div or slower) as well as during waveform record.
Return format	The query returns 1 or 0.
Example	:MASK:ENABle ON /*Enable the pass/fail test*/ / :MASK:ENABle? /*The query returns 1*/

## :MASK:SOURce
Syntax	:MASK:SOURce <source> / :MASK:SOURce?
Description	Set or query the source of the pass/fail test.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Explanation	Only channels enabled can be selected as the source. You can sent the :CHANnel<n>:DISPlay command to enable the desired channel.
Return format	The query returns CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:MASK:SOURce CHANnel2 /*Set the source of the pass/fail test to CH2*/ / :MASK:SOURce? /*The query returns CHAN2*/

## :MASK:OPERate
Syntax	:MASK:OPERate <oper> / :MASK:OPERate?
Description	Run or stop the pass/fail test, or query the status of the pass/fail test.
Parameter	Name	Type	Range	Default
	<oper>	Discrete	{RUN|STOP}	STOP
Explanation	Before executing this command, you need to send the :MASK:ENABle command to enable the pass/fail test.
Return format	The query returns RUN or STOP.
Example	:MASK:OPERate RUN /*Run the pass/fail test*/ / :MASK:OPERate? /*The query returns RUN*/

## :MASK:MDISplay
Syntax	:MASK:MDISplay <bool> / :MASK:MDISplay?
Description	Enable or disable the statistic information when the pass/fail test is enabled, or query the status of the statistic information.
Parameter	Name	Type	Range	Default
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Explanation	Before executing this command, you need to send the :MASK:ENABle command to enable the pass/fail test. / When the statistic information is on, the test results as shown in the figure below will be displayed at the upper right corner of the screen. / You can sent the :MASK:PASSed?, :MASK:FAILed? and :MASK:TOTal? commands to query the test results.
Return format	The query returns 1 or 0.
Example	:MASK:MDISplay ON /*Enable the statistic information*/ / :MASK:MDISplay? /*The query returns 1*/

## :MASK:SOOutput
Syntax	:MASK:SOOutput <bool> / :MASK:SOOutput?
Description	Turn the “Stop on Fail” function on or off, or query the status of the “Stop on Fail” function.
Parameter	Name	Type	Range	Default
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Explanation	ON: when failed waveforms are detected, the oscilloscope will stop the test and enter the “STOP” state. At this point, the results of the test remain the same on the screen (if the display is turned on) and the [Trigger Out] connector (if enabled) at the rear panel outputs a single pulse. / OFF: the oscilloscope will continue with the test even though failed waveforms are detected. The test results on the screen will update continuously and the [Trigger Out] connector at the rear panel outputs a pulse each time a failed waveform is detected.
Return format	The query returns 1 or 0.
Example	:MASK:SOOutput ON /*Turn the “Stop on Fail” function on*/ / :MASK:SOOutput? /*The query returns 1*/

## :MASK:OUTPut
Syntax	:MASK:OUTPut <bool> / :MASK:OUTPut?
Description	Enable or disable the sound prompt when failed waveforms are detected, or query the status of the sound prompt.
Parameter	Name	Type	Range	Default
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Explanation	OFF: when failed waveforms are detected, there are display and output but the beeper does not sound. / ON: when failed waveforms are detected, there are display and output and the beeper sounds (not related to the on/off state of the sound).
Return format	The query returns 1 or 0.
Example	:MASK:OUTPut ON /*Enable the sound prompt*/ / :MASK:OUTPut? /*The query returns 1*/

## :MASK:X
Syntax	:MASK:X <x> / :MASK:X?
Description	Set or query the horizontal adjustment parameter in the pass/fail test mask.
Parameter	Name	Type	Range	Default
	<x>	Real	0.02 to 4, the step is 0.02 within the range	0.24
Syntax	The query returns the horizontal adjustment parameter in scientific notation.
Description	:MASK:X 0.28 /*Set the horizontal adjustment parameter to 0.28div*/ / :MASK:X? /*The query returns 2.800000e-01*/

## :MASK:Y
Syntax	:MASK:Y <y> / :MASK:Y?
Description	Set or query the vertical adjustment parameter in the pass/fail test mask.
Parameter	Name	Type	Range	Default
	<y>	Real	0.04 to 5.12, the step is 0.04 within the range	0.48
Return format	The query returns the vertical adjustment parameter in scientific notation.
Example	:MASK:Y 0.36 /*Set the vertical adjustment parameter to 0.36div*/ / :MASK:Y? /*The query returns 3.600000e-01*/

## :MASK:CREate
Syntax	:MASK:CREate
Description	Create the pass/fail test mask using the current horizontal adjustment parameter and vertical adjustment parameter.
Explanation	This command is valid only when the pass/fail test is enabled (:MASK:ENABle) and is not in the run state (:MASK:OPERate).
Related commands	:MASK:X / :MASK:Y

## :MASK:PASSed?
Syntax	:MASK:PASSed?
Description	Query the number of passed frames in the pass/fail test.
Return Format	The query returns an integer.

## :MASK:FAILed?
Syntax	:MASK:FAILed?
Description	Query the number of failed frames in the pass/fail test.
Return Format	The query returns an integer.

## :MASK:TOTal?
Syntax	:MASK:TOTal?
Description	Query the total number of frames in the pass/fail test.
Return Format	The query returns an integer.

## :MASK:RESet
Syntax	:MASK:RESet
Description	Reset the numbers of passed frames and failed frames as well as the total number of frames in the pass/fail test to 0.
Related commands	:MASK:PASSed? / :MASK:FAILed? / :MASK:TOTal?