# :TRACe Commands (Only for MSO1000Z-S/DS1000Z-S)

The [:TRACe[<n>]] commands are used to set the arbitrary waveform parameters of the built-in signal sources. <n> can be 1 or 2 which denotes the corresponding built-in signal source channel. If <n> or :TRACe[<n>] is omitted, the operation will be applied to source 1 by default.

Command List:
- [:TRACe[<n>]]:DATA
- [:TRACe[<n>]]:DATA:DAC16
- [:TRACe[<n>]]:DATA:DAC
- [:TRACe[<n>]]:DATA:POINts
- [:TRACe[<n>]]:DATA:POINts:INTerpolate
- [:TRACe[<n>]]:DATA:VALue
- [:TRACe[<n>]]:DATA:LOAD?
## [:TRACe[<n>]]:DATA
Syntax	[:TRACe[<n>]]:DATA volatile,<value>[,<value>...]
Description	Download the floating point voltage values to the volatile memory of the specified signal source.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2}	1
	<value>	Real	-1 to 1	--
Explanation	<value> denotes the floating point voltage values. -1 and 1 correspond to the minimum and maximum values of the waveform amplitude respectively. For example, when the amplitude is 5Vpp and the offset is 0VDC, -1 corresponds to -2.5V and 1 corresponds to 2.5V. This command will overwrite the previous waveform in the volatile memory (no error will be generated). / 2 to 16384 points can be downloaded each time. / After sending this command, the specified signal source channel switches to volatile waveform output automatically. While at the same time, the initial number of points is modified. Users can edit the data downloaded using this command on the instrument.
Example	:DATA volatile,-0.5,-0.25,0.25,0.75 /*Download 4 floating point voltage values (-0.5,-0.25,0.25,0.75) to the volatile memory of the source 1*/

## [:TRACe[<n>]]:DATA:DAC16
Syntax	[:TRACe[<n>]]:DATA:DAC16 volatile,<flag>,<binary_block_data>
Description	Download binary data block to the volatile memory of the specified signal source.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2}	1
	<flag>	Discrete	{END}	--
	<binary_block_data>	Refer to the Explanation
Explanation	This command consists of two parts: the command string ([:TRACe[<n>]]:DATA:DAC16 volatile,<flag>,) and the binary data (<binary_block_data>). / <flag> denotes the data transmission status and can only be set to END which denotes the data transmission finishes. / <binary_block_data> denotes the binary data to be downloaded. The data length ranges from 4Bytes (2pt) to 32kBytes (16kpts). <binary_block_data> is a binary data block starts with #. For example, #516384 binary data; the number 5 behind # denotes that the data length information (16384) occupies 5 characters. 16384 denotes the number of bytes of the binary data. As each waveform point corresponds to two bytes of binary number (the range is from 0000 to 3FFF; 0000 and 3FFF correspond to the minimum and maximum values of the waveform amplitude respectively), the number of bytes must be an even number. / When END (data transmission complete indicator) is recieved, the instrument switches to arbitrary waveform output automatically.

## [:TRACe[<n>]]:DATA:DAC
Syntax	[:TRACe[<n>]]:DATA:DAC volatile,[<binary_block_data>|<value>,<value>,<value>...]
Description	Download binary data block or decimal waveform values to the volatile memory of the specified signal source.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2}	1
	<binary_block_data>	Refer to the Explanation
	<value>	Integer	0 to 16383	--
Explanation	<binary_block_data> denotes the binary data to be downloaded. The data length is from 4 Bytes (2 pts) to 32768 Bytes (16 kpts). <binary_block_data> is a binary data block starts with #. For example, #516384 binary data; the number 5 behind # denotes that the data length information (16384) occupies 5 characters. 16384 denotes the number of bytes of the binary data. As each waveform point corresponds to two bytes of binary number (the range is from 0000 to 3FFF; 0000 and 3FFF correspond to the minimum and maximum values of the waveform amplitude respectively), the number of bytes must be an even number. / <value> denotes the decimal DAC values. 0 and 16383 correspond to the minimum and maximum values of the waveform amplitude respectively. For example, when the amplitude is 5Vpp and the offset is 0VDC, 0 corresponds to -2.5V and 16383 corresponds to 2.5V. The range of the number of waveform points is from 2pts to 16384pts. For example, 5 waveform point data are sent by the :DATA:DAC volatile,0,16383,8192,0,16383 command. / After sending this command, the specified signal source channel switches to volatile waveform output automatically. While at the same time, the initial number of points is modified. Users can edit the data downloaded using this command on the instrument.

## [:TRACe[<n>]]:DATA:POINts
Syntax	[:TRACe[<n>]]:DATA:POINts volatile,<data> / [:TRACe[<n>]]:DATA:POINts? volatile
Description	Set or query the initial number of points of the arbitrary waveform of the specified signal source channel.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2}	1
	<data>	Integer	1 to 16384	2
Explanation	This command initializes the voltage value of each point on the volatile waveform to 0μV automatically. / After sending this command, you can use the [:TRACe[<n>]]:DATA:VALue command to modify the voltage of the specified point.
Return format	The query returns an integer between 1 and 16384.
Example	:DATA:POINts volatile,512 /*Set the initial number of points of the source 1 arbitrary waveform to 512*/ / :DATA:POINts? volatile /*The query returns 512*/

## [:TRACe[<n>]]:DATA:POINts:INTerpolate
Syntax	[:TRACe[<n>]]:DATA:POINts:INTerpolate <mode> [:TRACe[<n>]]:DATA:POINts:INTerpolate?
Description	Set or query the interpolation mode of the editable points of the arbitrary waveform of the specified signal source channel.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2}	1
	<mode>	Discrete	{LINear|OFF}	OFF
Explanation	LINear: turn on the linear interpolation. The waveform editor connects two editable points using a straight line. / OFF: turn off the linear interpolation. The waveform editor will keep a constant voltage level between two editable points and create a ladder-like waveform.
Return format	The query returns LIN or OFF.
Example	:DATA:POINts:INTerpolate LIN /*Set the interpolation mode of the source 1 arbitrary waveform to linear*/ / :DATA:POINts:INTerpolate? /*The query returns LIN*/

## [:TRACe[<n>]]:DATA:VALue
Syntax	[:TRACe[<n>]]:DATA:VALue volatile,<points>,<data> / [:TRACe[<n>]]:DATA:VALue? volatile,<points>
Description	Modify or query the decimal value of the specified point in the volatile memory of the specified signal source.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2}	1
	<points>	Integer	1 to the initial number of points	--
	<data>	Integer	0 to 16383	--
Explanation	This command is only valid when the current output is volatile waveform. / The initial number of points can be set by the [:TRACe[<n>]]:DATA:POINts command. / <data> denotes the decimal value. 0 and 16383 correspond to the minimum and maximum values of the waveform amplitude respectively. For example, when the amplitude is 5Vpp and the offset is 0VDC, 0 corresponds to -2.5V and 16383 corresponds to 2.5V.
Return format	The query returns an integer between 0 and 16383.
Example	:DATA:VALue volatile,1,10 /*Set the first point in the source 1 volatile memory to 10*/ / :DATA:VALue? volatile,1 /*The query returns 10*/

## [:TRACe[<n>]]:DATA:LOAD?
Syntax	[:TRACe[<n>]]:DATA:LOAD? <num>
Description	Read the specified data packet in the volatile memory of the specified signal source.
Parameter	Name	Type	Range	Default
	<n>	Discrete	{1|2}	1
	<num>	Integer	1	--
Return format	The query returns a binary data block. The data block header is #9000032768 followed by 32768 byte binary data.