# :TIMebase Commands

The :TIMebase commands are used to set the horizontal parameters, such as enabling the delayed sweep and setting the horizontal timebase mode.

Command List:
- :TIMebase:DELay:ENABle
- :TIMebase:DELay:OFFSet
- :TIMebase:DELay:SCALe
- :TIMebase[:MAIN]:OFFSet
- :TIMebase[:MAIN]:SCALe
- :TIMebase:MODE
## :TIMebase:DELay:ENABle
Syntax	:TIMebase:DELay:ENABle <bool> / :TIMebase:DELay:ENABle?
Description	Enable or disable the delayed sweep, or query the status of the delayed sweep.
Parameter	Name	Type	Range	Default
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Explanation	Delayed sweep can be used to enlarge a length of waveform horizontally to view waveform details.
Return format	The query returns 1 or 0.
Example	:TIMebase:DELay:ENABle ON /*Enable the delayed sweep*/ / :TIMebase:DELay:ENABle? /*The query returns 1*/

## :TIMebase:DELay:OFFSet
Syntax	:TIMebase:DELay:OFFSet <offset> / :TIMebase:DELay:OFFSet?
Description	Set or query the delayed timebase offset. The default unit is s.
Parameter	Name	Type	Range	Default
	<offset>	Real	-(LeftTime - DelayRange/2) to / (RightTime - DelayRange/2)	0
Explanation	LeftTime = 6×MainScale – MainOffset / RightTime = 6×MainScale + MainOffset DelayRange = 12×DelayScale
Return format	The query returns the offset in scientific notation.
Example	:TIMebase:DELay:OFFSet 0.000002 /*Set the delayed timebase offset to 2μs*/ :TIMebase:DELay:OFFSet? /*The query returns 2.0000000e-06*/
Related commands	:TIMebase:DELay:SCALe / :TIMebase[:MAIN]:OFFSet / :TIMebase[:MAIN]:SCALe

## :TIMebase:DELay:SCALe
Syntax	:TIMebase:DELay:SCALe <scale_value> / :TIMebase:DELay:SCALe?
Description	Set or query the delayed timebase scale. The default unit is s/div.
Parameter	Name	Type	Range	Default
	<scale_value>	Real	See Explanation	500ns
Explanation	The maximum of <scale_value> is the main timebase scale currently set and the minimum is related to the current sample rate (satisfying the inequality: current sample rate * delayed timebase scale * 7 ≥ 14) and is no less than 1ns. / The delayed timebase scale can only be the maximum or the value decreased from the maximum in 1-2-5 step. If the minimum calculated by the inequality mentioned above is not a settable value, the larger settable value that is nearest to the calculated value will be used as the minimum. If the value set is between two settable values, the smaller one will be used as the delayed timebase scale automatically. / For example, when the main timebase scale currently set is 10.00μs and the sample rate is 500MSa/s, the maximum of <scale_value> is 10.00μs. The minimum is calculated using the inequality: ≥ 14 / (500MSa/s * 7) = 40ns. Taking the larger settable value that is nearest to the calculated value, the minimum is 50ns. At this point, if 0.000009 (namely 9μs) is set, the instrument sets the delayed timebase scale to 5μs.
Return format	The query returns the scale in scientific notation.
Example	:TIMebase:DELay:SCALe 0.00000005 /*Set the delayed timebase scale to 50ns*/ / :TIMebase:DELay:SCALe? /*The query returns 5.0000000e-08*/
Related command	:TIMebase[:MAIN]:SCALe

## :TIMebase[:MAIN]:OFFSet
Syntax	:TIMebase[:MAIN]:OFFSet <offset> / :TIMebase[:MAIN]:OFFSet?
Description	Set or query the main timebase offset. The default unit is s.
Parameter	Name	Type	Range	Default
	<offset>	Real	-Screen/2 to 1s or -Screen/2 to 5000s	0
Explanation	This command is invalid when the horizontal timebase mode is Roll and the oscilloscope is in the Run state. / When the horizontal timebase mode is YT and the horizontal timebase is 200ms or larger (namely the “Slow Sweep” mode), this command is invalid when the oscilloscope in the Stoping state.
Return format	The query returns the offset in scientific notation.
Example	:TIMebase:MAIN:OFFSet 0.0002 /*Set the main timebase offset to 20ms*/ / :TIMebase:MAIN:OFFSet? /*The query returns 2.0000000e-04*/
Related commands	:TIMebase:MODE / :TIMebase[:MAIN]:SCALe / :RUN / :STOP

## :TIMebase[:MAIN]:SCALe
Syntax	:TIMebase[:MAIN]:SCALe <scale_value> / :TIMebase[:MAIN]:SCALe?
Description	Set or query the main timebase scale. The default unit is s/div.
Parameter	Name	Type	Range	Default
	<scale_value>	Real	Normal: 5ns to 50s in 1-2-5 step Roll: 200ms to 50s in 1-2-5 step	1μs
Explanation	When the horizontal timebase mode is YT and the horizontal timebase is 200ms or larger (namely the “Slow Sweep” mode), this command is invalid when the oscilloscope in the Stoping state.
Return format	The query returns the main timebase scale in scientific notation.
Example	:TIMebase:MAIN:SCALe 0.0002 /*Set the main timebase scale to 200μs*/ / :TIMebase:MAIN:SCALe? /*The query returns 2.0000000e-04*/
Related command	:TIMebase:MODE

## :TIMebase:MODE
Syntax	:TIMebase:MODE <mode> / :TIMebase:MODE?
Description	Set or query the mode of the horizontal timebase.
Parameter	Name	Type	Range	Default
	<mode>	Discrete	{MAIN|XY|ROLL}	MAIN
Return format	The query returns MAIN, XY or ROLL.
Example	:TIMebase:MODE XY /*Set the horizontal timebase mode to XY*/ / :TIMebase:MODE? /*The query returns XY*/
Related commands	:TIMebase:DELay:OFFSet / :TIMebase:DELay:SCALe / :TIMebase[:MAIN]:OFFSet / :TIMebase[:MAIN]:SCALe