# :TRIGger Commands

The :TRIGger commands are used to set the trigger system of the oscilloscope.

Command List:
- :TRIGger:MODE
- :TRIGger:COUPling
- :TRIGger:STATus?
- :TRIGger:SWEep
- :TRIGger:HOLDoff
- :TRIGger:NREJect
- :TRIGger:EDGe
- :TRIGger:PULSe
- :TRIGger:SLOPe
- :TRIGger:VIDeo
- :TRIGger:PATTern
- :TRIGger:DURATion
- :TRIGger:TIMeout (Option)
- :TRIGger:RUNT (Option)
- :TRIGger:WINDows (Option)
- :TRIGger:DELay (Option)
- :TRIGger:SHOLd (Option)
- :TRIGger:NEDGe (Option)
- :TRIGger:RS232 (Option)
- :TRIGger:IIC (Option)
- :TRIGger:SPI (Option)
## :TRIGger:MODE
Syntax	:TRIGger:MODE <mode> / :TRIGger:MODE?
Description	Select or query the trigger type.
Parameter	Name	Type	Range	Default
	<mode>	Discrete	{EDGE|PULSe|RUNT|WIND|NEDG|SLOPe|VIDeo| / PATTern|DELay|TIMeout|DURation|SHOLd|RS232 |IIC|SPI}	EDGE
Explanation
Return format	The query returns EDGE, PULS, RUNT, WIND, NEDG, SLOP, VID, PATT, DEL, TIM, DUR, SHOL, RS232, IIC or SPI.
Example	:TRIGger:MODE SLOPe /*Select slope trigger*/ / :TRIGger:MODE? /*The query returns SLOP*/

## :TRIGger:COUPling
Syntax	:TRIGger:COUPling <couple> / :TRIGger:COUPling?
Description	Select or query the trigger coupling type.
Parameter	Name	Type	Range	Default
	<couple>	Discrete	{AC|DC|LFReject|HFReject}	DC
Explanation	This command is only applicable to edge trigger of which the signal source is an analog channel. / AC: block all the DC components and attenuate signals lower than 75 kHz. / DC: allow DC and AC components into the trigger path. / LFReject: block the DC components and reject the low frequency components (lower than 75 kHz). / HFReject: reject the high frequency components (higher than 75 kHz).
Return format	The query returns AC, DC, LFR or HFR.
Example	:TRIGger:COUPling LFReject /*Set the trigger coupling type to low-frequency rejection*/ / :TRIGger:COUPling? /*The query returns LFR*/

## :TRIGger:STATus?
Syntax	:TRIGger:STATus?
Description	Query the current trigger status.
Parameter	The query returns TD, WAIT, RUN, AUTO or STOP.

## :TRIGger:SWEep
Syntax	:TRIGger:SWEep <sweep> / :TRIGger:SWEep?
Description	Set or query the trigger mode.
Parameter	Name	Type	Range	Default
	<sweep>	Discrete	{AUTO|NORMal|SINGle}	AUTO
Explanation	AUTO: auto trigger. No matter whether the trigger condition is met, there is always waveform display. / NORMal: normal trigger. Display waveform when the trigger condition is met; otherwise, the oscilloscope holds the original waveform and waits for the next trigger. / SINGle: single trigger. The oscilloscope waits for a trigger and displays the waveform when the trigger condition is met and then stops.
Return format	The query returns AUTO, NORM or SING.
Example	:TRIGger:SWEep SINGle /*Select single trigger mode */ / :TRIGger:SWEep? /*The query returns SING*/

## :TRIGger:HOLDoff
Syntax	:TRIGger:HOLDoff <value> / :TRIGger:HOLDoff?
Description	Set or query the trigger holdoff time. The default unit is s.
Parameter	Name	Type	Range	Default
	<value>	Real	16ns to 10s	16ns
Explanation	Trigger holdoff can be used to stably trigger the complex waveforms (such as pulse series). Holdoff time is the time that the oscilloscope waits before re-arming the trigger circuitry. The oscilloscope will not trigger until the holdoff time expires. / When the trigger type is video, timeout, setup/hold, Nth edge, RS232, I2C or SPI, this setting will be not available.
Return format	The query returns the trigger holdoff time in scientific notation.
Example	:TRIGger:HOLDoff 0.0000002 /*Set the trigger holdoff time to 200ns*/ / :TRIGger:HOLDoff? /*The query returns 2.000000e-07*/

## :TRIGger:NREJect
Syntax	:TRIGger:NREJect <bool> / :TRIGger:NREJect?
Description	Enable or disable noise rejection, or query the status of noise rejection.
Parameter	Name	Type	Range	Default
	<bool>	Bool	{{1|ON}|{0|OFF}}	0|OFF
Explanation	Noise rejection reduces the possibility of noise trigger. / This command is not available when the trigger source is a digital channel.
Return format	The query returns 1 or 0.
Example	:TRIGger:NREJect ON /*Enable noise rejection*/ / :TRIGger:NREJect? /*The query returns 1*/

# :TRIGger:EDGe

Command List:
- :TRIGger:EDGe:SOURce
- :TRIGger:EDGe:SLOPe
- :TRIGger:EDGe:LEVel
## :TRIGger:EDGe:SOURce
Syntax	:TRIGger:EDGe:SOURce <source> / :TRIGger:EDGe:SOURce?
Description	Set or query the trigger source in edge trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4|AC}	CHANnel1
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3, CHAN4 or AC.
Example	:TRIGger:EDGe:SOURce CHANnel1 /*Set the trigger source to CH1*/ / :TRIGger:EDGe:SOURce? /*The query returns CHAN1*/

## :TRIGger:EDGe:SLOPe
Syntax	:TRIGger:EDGe:SLOPe <slope> / :TRIGger:EDGe:SLOPe?
Description	Set or query the edge type in edge trigger.
Parameter	Name	Type	Range	Default
	<slope>	Discrete	{POSitive|NEGative|RFALl}	POSitive
Explanation	POSitive: rising edge / NEGative: falling edge / RFALl: rising/falling edge
Return format	The query returns POS, NEG or RFAL.
Example	:TRIGger:EDGe:SLOPe NEGative /*Set the edge type to falling edge*/ / :TRIGger:EDGe:SLOPe? /*The query returns NEG*/
Related command	:TRIGger:MODE

## :TRIGger:EDGe:LEVel
Syntax	:TRIGger:EDGe:LEVel <level> / :TRIGger:EDGe:LEVel?
Description	Set or query the trigger level in edge trigger. The unit is the same as the current amplitude unit of the signal source selected.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - OFFSet	0
Explanation	For VerticalScale, refer to the :CHANnel<n>:SCALe command. For OFFSet, refer to the :CHANnel<n>:OFFSet command. / This command is only available when the signal source is an analog channel.
Return format	The query returns the trigger level in scientific notation.
Example	:TRIGger:EDGe:LEVel 0.16 /*Set the trigger level to 160mV*/ / :TRIGger:EDGe:LEVel? /*The query returns 1.600000e-01*/

# :TRIGger:PULSe

Command List:
- :TRIGger:PULSe:SOURce
- :TRIGger:PULSe:WHEN
- :TRIGger:PULSe:WIDTh
- :TRIGger:PULSe:UWIDth
- :TRIGger:PULSe:LWIDth
- :TRIGger:PULSe:LEVel
## :TRIGger:PULSe:SOURce
Syntax	:TRIGger:PULSe:SOURce <source> / :TRIGger:PULSe:SOURce?
Description	Set or query the trigger source in pulse width trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| / D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:PULSe:SOURce CHANnel1 /*Set the trigger source to CH1*/ / :TRIGger:PULSe:SOURce? /*The query returns CHAN1*/

## :TRIGger:PULSe:WHEN
Syntax	:TRIGger:PULSe:WHEN <when> / :TRIGger:PULSe:WHEN?
Description	Set or query the trigger condition in pulse width trigger.
Parameter	Name	Type	Range	Default
	<when>	Discrete	{PGReater|PLESs|NGReater| / NLESs|PGLess|NGLess}	PGReater
Explanation	PGReater/PLESs: you need to specify a pulse width (refer to :TRIGger:PULSe:WIDTh). The oscilloscope triggers when the positive pulse width of the input signal is greater/lower than the specified Pulse Width. / NGReater/NLESs: you need to specify a pulse width (refer to :TRIGger:PULSe:WIDTh). The oscilloscope triggers when the negative pulse width of the input signal is greater/lower than the specified Pulse Width. / PGLess/NGLess: you need to specify an upper (refer to :TRIGger:PULSe:UWIDth) and a lower (refer to :TRIGger:PULSe:LWIDth) pulse width. The oscilloscope triggers when the positive/negative pulse width of the input signal is greater than the specified lower pulse width and lower than the upper pulse width.
Return format	The query returns PGR, PLES, NGR, NLES, PGL or NGL.
Example	:TRIGger:PULSe:WHEN NLESs /*Set the trigger condition to NLESs*/ / :TRIGger:PULSe:WHEN? /*The query returns NELS*/

## :TRIGger:PULSe:WIDTh
Syntax	:TRIGger:PULSe:WIDTh <width> / :TRIGger:PULSe:WIDTh?
Description	Set or query the pulse width in pulse width trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<width>	Real	8ns to 10s	PGReater, NGReater: 1μs / PLESs, NLESs: 2μs
Explanation	This command is available when the trigger condition (refer to :TRIGger:PULSe:WHEN) is PGReater, PLESs, NGReater and NLESs.
Return format	The query returns the pulse width in scientific notation.
Example	:TRIGger:PULSe:WIDTh 0.000003 /*Set the pulse width to 3μs*/ / :TRIGger:PULSe:WIDTh? /*The query returns3.000000e-06*/

## :TRIGger:PULSe:UWIDth
Syntax	:TRIGger:PULSe:UWIDth <width> / :TRIGger:PULSe:UWIDth?
Description	Set or query the upper pulse width in pulse width trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<width>	Real	16ns to 10s	2μs
Explanation	This command is available when the trigger condition (refer to :TRIGger:PULSe:WHEN) is PGLess and NGLess.
Return format	The query returns the upper pulse width in scientific notation.
Example	:TRIGger:PULSe:UWIDth 0.000003 /*Set the upper pulse width to 3μs*/ / :TRIGger:PULSe:UWIDth? /*The query returns3.000000e-06*/
Related command	:TRIGger:PULSe:LWIDth

## :TRIGger:PULSe:LWIDth
Syntax	:TRIGger:PULSe:LWIDth <width> / :TRIGger:PULSe:LWIDth?
Description	Set or query the lower pulse width in pulse width trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<width>	Real	8ns to 9.99s	1μs
Explanation	This command is available when the trigger condition (refer to :TRIGger:PULSe:WHEN) is PGLess and NGLess.
Return format	The query returns the lower pulse width in scientific notation.
Example	:TRIGger:PULSe:LWIDth 0.000003 /*Set the lower pulse width to 3μs*/ / :TRIGger:PULSe:LWIDth? /*The query returns 3.000000e-06*/
Related command	:TRIGger:PULSe:UWIDth

## :TRIGger:PULSe:LEVel
Syntax	:TRIGger:PULSe:LEVel <level> / :TRIGger:PULSe:LEVel?
Description	Set or query the trigger level in pulse width trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - / OFFSet	0
Explanation	For VerticalScale, refer to the :CHANnel<n>:SCALe command. For OFFSet, refer to the :CHANnel<n>:OFFSet command. / This command is only available when the signal source is an analog channel.
Return format	The query returns the trigger level in scientific notation.
Example	:TRIGger:PULSe:LEVel 0.16 /*Set the trigger level to 160mV*/ / :TRIGger:PULSe:LEVel? /*The query returns 1.600000e-01*/

# :TRIGger:SLOPe

Command List:
- :TRIGger:SLOPe:SOURce
- :TRIGger:SLOPe:WHEN
- :TRIGger:SLOPe:TIME
- :TRIGger:SLOPe:TUPPer
- :TRIGger:SLOPe:TLOWer
- :TRIGger:SLOPe:WINDow
- :TRIGger:SLOPe:ALEVel
- :TRIGger:SLOPe:BLEVel
## :TRIGger:SLOPe:SOURce
Syntax	:TRIGger:SLOPe:SOURce <source> / :TRIGger:SLOPe:SOURce?
Description	Set or query the trigger source in slope trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Return format	The query returns CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:SLOPe:SOURce CHANnel2 /*Set the trigger source to CH2*/ / :TRIGger:SLOPe:SOURce? /*The query returns CHAN2*/

## :TRIGger:SLOPe:WHEN
Syntax	:TRIGger:SLOPe:WHEN <when> / :TRIGger:SLOPe:WHEN?
Description	Set or query the trigger condition in slope trigger.
Parameter	Name	Type	Range	Default
	<when>	Discrete	{PGReater|PLESs|NGReater| / NLESs|PGLess|NGLess}	PGReater
Explanation	PGReater/PLESs: you need to specify a time value (refer to :TRIGger:SLOPe:TIME). The oscilloscope triggers when the positive slope time of the input signal is greater/lower than the specified time. / NGReater/NLESs: you need to specify a time value (refer to :TRIGger:SLOPe:TIME). The oscilloscope triggers when the negative slope time of the input signal is greater/lower than the specified time. / PGLess/NGLess: you need to specify an upper limit (refer to :TRIGger:SLOPe:TUPPer) and a lower limit (refer to :TRIGger:SLOPe:TLOWer) of the time. The oscilloscope triggers when the positive/negative slope time of the input signal is greater than the specified lower limit and lower than the specified upper limit.
Return format	The query returns PGR, PLES, NGR, NLES, PGL or NGL.
Example	:TRIGger:SLOPe:WHEN NLESs /*Set the trigger condition to NLESs*/ / :TRIGger:SLOPe:WHEN? /*The query returns NLES*/

## :TRIGger:SLOPe:TIME
Syntax	:TRIGger:SLOPe:TIME <time> / :TRIGger:SLOPe:TIME?
Description	Set or query the time value in slope trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<time>	Real	8ns to 10s	PGReater, NGReater: 1μs / PLESs, NLESs: 2μs
Explanation	This command is available when the trigger condition (refer to :TRIGger:SLOPe:WHEN) is PGReater, PLESs, NGReater and NLESs.
Return format	The query returns the time value in scientific notation.
Example	:TRIGger:SLOPe:TIME 0.000003 /*Set the time value to 3μs*/ / :TRIGger:SLOPe:TIME? /*The query returns 3.000000e-06*/

## :TRIGger:SLOPe:TUPPer
Syntax	:TRIGger:SLOPe:TUPPer <time> / :TRIGger:SLOPe:TUPPer?
Description	Set or query the upper limit of the time in slope trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<time>	Real	16ns to 10s	2μs
Explanation	This command is available when the trigger condition (refer to :TRIGger:SLOPe:WHEN) is PGLess and NGLess.
Return format	The query returns the upper limit of the time in scientific notation.
Example	:TRIGger:SLOPe:TUPPer 0.000003 /*Set the upper limit of the time to 3μs*/ / :TRIGger:SLOPe:TUPPer? /*The query returns 3.000000e-06*/
Related command	:TRIGger:SLOPe:TLOWer

## :TRIGger:SLOPe:TLOWer
Syntax	:TRIGger:SLOPe:TLOWer <time> / :TRIGger:SLOPe:TLOWer?
Description	Set or query the lower limit of the time in slope trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<time>	Real	8ns to 9.99s	1μs
Explanation	This command is available when the trigger condition (refer to :TRIGger:SLOPe:WHEN) is PGLess and NGLess.
Return format	The query returns the lower limit of the time in scientific notation.
Example	:TRIGger:SLOPe:TLOWer 0.000000020 /*Set the lower limit of the time to 20ns*/ / :TRIGger:SLOPe:TLOWer? /*The query returns 2.000000e-08*/
Related command	:TRIGger:SLOPe:TUPPer

## :TRIGger:SLOPe:WINDow
Syntax	:TRIGger:SLOPe:WINDow <window> / :TRIGger:SLOPe:WINDow?
Description	Set or query the vertical window type in slope trigger.
Parameter	Name	Type	Range	Default
	<window>	Discrete	{TA|TB|TAB}	TA
Explanation	TA: only adjust the upper limit of the trigger level, refer to :TRIGger:SLOPe:ALEVel. / TB: only adjust the lower limit of the trigger level, refer to :TRIGger:SLOPe:BLEVel. / TAB: adjust the upper and lower limits of the trigger level at the same time, refer to :TRIGger:SLOPe:ALEVel and :TRIGger:SLOPe:BLEVel.
Return format	The query returns TA, TB or TAB.
Example	:TRIGger:SLOPe:WINDow TB /*Set the vertical window type to TB*/ / :TRIGger:SLOPe:WINDow? /*The query returns TB*/

## :TRIGger:SLOPe:ALEVel
Syntax	:TRIGger:SLOPe:ALEVel <level> / :TRIGger:SLOPe:ALEVel?
Description	Set or query the upper limit of the trigger level in slope trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - / OFFSet	2V
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet.
Return format	The query returns the upper limit of the trigger level in scientific notation.
Example	:TRIGger:SLOPe:ALEVel 0.16 /*Set the upper limit of the trigger level to 160mV*/ / :TRIGger:SLOPe:ALEVel? /*The query returns 1.600000e-01*/
Related command	:TRIGger:SLOPe:BLEVel

## :TRIGger:SLOPe:BLEVel
Syntax	:TRIGger:SLOPe:BLEVel <level> / :TRIGger:SLOPe:BLEVel?
Description	Set or query the lower limit of the trigger level in slope trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet.
Return format	The query returns the lower limit of the trigger level in scientific notation.
Example	:TRIGger:SLOPe:BLEVel 0.16 /*Set the lower limit of the trigger level to 160mV*/ / :TRIGger:SLOPe:BLEVel? /*The query returns 1.600000e-01*/
Related command	:TRIGger:SLOPe:ALEVel

# :TRIGger:VIDeo

Command List:
- :TRIGger:VIDeo:SOURce
- :TRIGger:VIDeo:POLarity
- :TRIGger:VIDeo:MODE
- :TRIGger:VIDeo:LINE
- :TRIGger:VIDeo:STANdard
- :TRIGger:VIDeo:LEVel
## :TRIGger:VIDeo:SOURce
Syntax	:TRIGger:VIDeo:SOURce <source> / :TRIGger:VIDeo:SOURce?
Description	Select or query the trigger source in video trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Return format	The query returns CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:VIDeo:SOURce CHANnel2 /*Set the trigger source to CH2*/ :TRIGger:VIDeo:SOURce? /*The query returns CHAN2*/

## :TRIGger:VIDeo:POLarity
Syntax	:TRIGger:VIDeo:POLarity <polarity> / :TRIGger:VIDeo:POLarity?
Description	Select or query the video polarity in video trigger.
Parameter	Name	Type	Range	Default
	<polarity>	Discrete	{POSitive|NEGative}	POSitive
Return format	The query returns POS or NEG.
Example	:TRIGger:VIDeo:POLarity POSitive /*Set or query the video polarity to positive*/ / :TRIGger:VIDeo:POLarity? /*The query returns POS*/

## :TRIGger:VIDeo:MODE
Syntax	:TRIGger:VIDeo:MODE <mode> / :TRIGger:VIDeo:MODE?
Description	Set or query the sync type in video trigger.
Parameter	Name	Type	Range	Default
	<mode>	Discrete	{ODDField|EVENfield|LINE|ALINes}	ALINes
Explanation	ODDField: trigger on the rising edge of the first ramp waveform pulse in the odd field. This type is available when the video standard is NTSC or PAL/SECAM. / EVENfield: trigger on the rising edge of the first ramp waveform pulse in the even field. This type is available when the video standard is NTSC or PAL/SECAM. / LINE: for the NTSC and PAL/SECAM video standards, trigger on the specified line in the odd or even field; for the 480P and 576P video standards, trigger on the specified line. / ALINes: trigger on all the horizontal sync pulses.
Return format	The query returns ODDF, EVEN, LINE or ALIN.
Example	:TRIGger:VIDeo:MODE ODDField /*Set the sync type to odd field*/ / :TRIGger:VIDeo:MODE? /*The query returns ODDF*/
Related commands	:TRIGger:VIDeo:LINE / :TRIGger:VIDeo:STANdard

## :TRIGger:VIDeo:LINE
Syntax	:TRIGger:VIDeo:LINE <line> / :TRIGger:VIDeo:LINE?
Description	Set or query the line number when the sync type in video trigger is LINE.
Parameter	Name	Type	Range	Default
	<line>	Integer	Refer to Explanation	1
Explanation	NTSC: 1 to 525 / PAL/SECAM: 1 to 625 / 480P: 1 to 525 / 576P: 1 to 625
Return format	The query returns an integer.
Example	:TRIGger:VIDeo:LINE 100 /*Set the line number to 100*/ / :TRIGger:VIDeo:LINE? /*The query returns 100*/
Related command	:TRIGger:VIDeo:MODE

## :TRIGger:VIDeo:STANdard
Syntax	:TRIGger:VIDeo:STANdard <standard> / :TRIGger:VIDeo:STANdard?
Description	Set or query the video standard in video trigger.
Parameter	Name	Type	Range	Default
	<standard>	Discrete	{PALSecam|NTSC|480P|576P}	NTSC
Explanation	PALSecam: / PAL: the frame frequency is 25 frames per second. The TV sweep line is 625 with the odd field goes first and the even field follows behind. / Secam: the frame frequency is 25 frames per second. The sweep line is 625 with interlacing sweep. / NTSC: the field frequency is 60 fields per second and the frame frequency is 30 frames per second. The TV sweep line is 525 with the even field goes first and the odd field follows behind. / 480P: the frame frequency is 60 frames per second; the TV sweep line is 525; line-by-line sweep; the line frequency is 31.5 kHz. / 576P: the frame frequency is 60 frames per second; the TV sweep line is 625; line-by-line sweep.
Return format	The query returns PALS, NTSC, 480P or 576P.
Example	:TRIGger:VIDeo:STANdard NTSC /*Select NTSC video standard*/ / :TRIGger:VIDeo:STANdard? /*The query returns NTSC*/
Related commands	:TRIGger:VIDeo:LINE / :TRIGger:VIDeo:MODE

## :TRIGger:VIDeo:LEVel
Syntax	:TRIGger:VIDeo:LEVel <level> / :TRIGger:VIDeo:LEVel?
Description	Set or query the trigger level in video trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - / OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. / For OFFSet, refer to :CHANnel<n>:OFFSet.
Return format	The query returns the trigger level in scientific notation.
Example	:TRIGger:VIDeo:LEVel 0.16 /*Set the trigger level to 160mV*/ / :TRIGger:VIDeo:LEVel? /*The query returns 1.600000e-01*/

# :TRIGger:PATTern

Command List:
- :TRIGger:PATTern:PATTern
- :TRIGger:PATTern:LEVel
## :TRIGger:PATTern:PATTern
Syntax	:TRIGger:PATTern:PATTern <pa_ch1>[,<pa_ch2>[,<pa_ch3> [,<pa_ch4>[,<pa_d0>…[,<pa_d15>]]]]] / :TRIGger:PATTern:PATTern?
Description	Set or query the pattern of each channel in pattern trigger.
Parameter	Name	Type	Range	Default
	<pa_ch1>	Discrete	{H|L|X|R|F}	X
	<pa_ch2>	Discrete	{H|L|X|R|F}	X
	<pa_ch3>	Discrete	{H|L|X|R|F}	X
	<pa_ch4>	Discrete	{H|L|X|R|F}	X
	<pa_d0>	Discrete	{H|L|X|R|F}	X
	……	……	……	……
	<pa_d15>	Discrete	{H|L|X|R|F}	X
	<pa_ch1> to <pa_ch4> set the patterns of analog channels CH1 to CH4. <pa_d0> to <pa_d15> set the patterns of digital channels D0 to D15. / Users can send 20 parameters to set the pataterns of all the channels. Users can also omit some parameters to set the patterns of some of the channels (the pattern states of the channels of which the parameters are omitted remains unchanged), but at least one parameter (this parameter sets the pattern of CH1) should be sent. When the number of parameters sent is less than 20, the instrument sets the channels in CH1 to CH4 and D0 to D15 order. / In the range of the parameter, H represents high level (higher than the threshold level of the channel). L represents low level (lower than the threshold level of the channel). X denotes that this channel is ignored (this channel is not used as a part of the pattern. When all the channels in the pattern are set to X, the oscilloscope will not trigger). R represents the rising edge and F represents the falling edge (in the pattern, only a single rising edge or falling edge can be defined. If an edge item is currently defined and then another edge item is defined in another channel in the pattern, the former edge item defined will be replaced by X). / Please distinguish “omit parameter” from “ignore channel”. The former means that the parameters corresponding to some channels are not sent when sending the command and the pattern states of these channels remain unchanged. The latter means that the channel is set to X when sending the command and the state of this channel is ignored when the instrument judges the patterns. / The query returns the patterns of the 4 analog channels when all the digital channels are turned off; the query returns the patterns of all the 20 channels when any of the digital channels is turned on.
Explanation	The query returns the patterns of the 4 analog channels or all the channels. The patterns of multiple channels are separated by commas.
Return format	:TRIGger:PATTern:PATTern H,R,L,X /*Set the patterns of CH1 to CH4 to H,R,L,X. The / patterns of the other channels remain unchanged*/ / :TRIGger:PATTern:PATTern? /*The query returns / H,R,L,X,X,X,X,X,X,X,X,X,X,X,X,X,X,X,X,X*/

## :TRIGger:PATTern:LEVel
Syntax	:TRIGger:PATTern:LEVel <chan>,<level> / :TRIGger:PATTern:LEVel? <chan>
Description	Set or query the trigger level of the specified channel in pattern trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<chan>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Explanation	<level>	Real	(± 5 × VerticalScale from the screen center) - / OFFSet	0
Return format	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet. / This command is only available when the signal source is an analog channel.
Example	The query returns the trigger level in scientific notation.
Syntax	:TRIGger:PATTern:LEVel CHANnel2,0.16 /*Set the trigger level of CH2 to 160mV*/ / :TRIGger:PATTern:LEVel? CHANnel2 /*The query returns 1.600000e-01*/

# :TRIGger:DURATion

Command List:
- :TRIGger:DURATion:SOURce
- :TRIGger:DURATion:TYPe
- :TRIGger:DURATion:WHEN
- :TRIGger:DURATion:TUPPer
- :TRIGger:DURATion:TLOWer
## :TRIGger:DURATion:SOURce
Syntax	:TRIGger:DURATion:SOURce <source> / :TRIGger:DURATion:SOURce?
Description	Set or query the trigger source in duration trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| / D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:DURATion:SOURce CHANnel2 /*Set the trigger source to CH2*/ / :TRIGger:DURATion:SOURce? /*The query returns CHAN2*/

## :TRIGger:DURATion:TYPe
Syntax	:TRIGger:DURATion:TYPe <type_ch1>[,<type_ch2>[,<type_ch3> [,<type_ch4>[,<type_d0>…[,<type_d15>]]]]] / :TRIGger:DURATion:TYPe?
Description	Set or query the pattern of each channel in duration trigger.
Parameter	Name	Type	Range	Default
	<type_ch1>	Discrete	{H|L|X}	X
	<type_ch2>	Discrete	{H|L|X}	X
	<type_ch3>	Discrete	{H|L|X}	X
	<type_ch4>	Discrete	{H|L|X}	X
	<type_d0>	Discrete	{H|L|X}	X
	……	……	……	……
	<type_d15>	Discrete	{H|L|X}	X
Explanation	<type_ch1> to <type_ch4> set the patterns of analog channels CH1 to CH4. <type_d0> to <type_d15> set the patterns of digital channels D0 to D15. / Users can send 20 parameters to set the pataterns of all the channels. Users can also omit some parameters to set the patterns of some of the channels (the pattern states of the channels of which the parameters are omitted remains unchanged), but at least one parameter (this parameter sets the pattern of CH1) should be sent. When the number of parameters sent is less than 20, the instrument sets the channels in CH1 to CH4 and D0 to D15 order. / In the range of the parameter, H represents high level (higher than the threshold level of the channel). L represents low level (lower than the threshold level of the channel). X denotes that this channel is ignored (this channel is not used as a part of the pattern. When all the channels in the pattern are set to X, the oscilloscope will not trigger). / Please distinguish “omit parameter” from “ignore channel”. The former means that the parameters corresponding to some channels are not sent when sending the command and the pattern states of these channels remain unchanged. The latter means that the channel is set to X when sending the command and the state of this channel is ignored when the instrument judges the patterns. / The query returns the patterns of the 4 analog channels when all the digital channels are turned off; the query returns the patterns of all the 20 channels when any of the digital channels is turned on.
Return format	The query returns the patterns of the 4 analog channels or all the channels. The patterns of multiple channels are separated by commas.
Example	:TRIGger:DURATion:TYPe L,X,H,L /*Set the patterns of CH1 to CH4 to L,X,H,L. The patterns of the other channels remain unchanged*/ / :TRIGger:DURATion:TYPe? /*The query returns / L,X,H,L,X,X,X,X,X,X,X,X,X,X,X,X,X,X,X,X*/

## :TRIGger:DURATion:WHEN
Syntax	:TRIGger:DURATion:WHEN <when> / :TRIGger:DURATion:WHEN?
Description	Set or query the trigger condition in duration trigger.
Parameter	Name	Type	Range	Default
	<when>	Discrete	{GREater|LESS|GLESs}	GREater
Explanation	GREater: you need to specify a time (refer to :TRIGger:DURATion:TLOWer). The oscilloscope triggers when the duration of the pattern is greater than the preset time. / LESS: you need to specify a time (refer to :TRIGger:DURATion:TUPPer). The oscilloscope triggers when the duration of the pattern is lower than the preset time. / GLESs: you need to specify an upper limit of the time (refer to :TRIGger:DURATion:TUPPer) and lower limit of the time (refer to :TRIGger:DURATion:TLOWer). The oscilloscope triggers when the duration of the pattern is lower than the preset upper limit of the time and greater than the preset lower limit of the time.
Return format	The query returns GRE, LESS or GLES.
Example	:TRIGger:DURATion:WHEN LESS /*Set the trigger condition to <*/ / :TRIGger:DURATion:WHEN? /*The query returns LESS*/

## :TRIGger:DURATion:TUPPer
Syntax	:TRIGger:DURATion:TUPPer <NR3> / :TRIGger:DURATion:TUPPer?
Description	Set or query the duration time upper limit in duration trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Real	Related to the trigger condition LESS: 8ns to 10s / GLESs: 16ns to 10s	2μs
Explanation	This command is available when the trigger condition (:TRIGger:DURATion:WHEN) is LESS or GLESs.
Return format	The query returns the duration time upper limit in scientific notation.
Example	:TRIGger:DURATion:TUPPer 0.000003 /*Set the duration time upper limit to 3μs*/ / :TRIGger:DURATion:TUPPer? /*The query returns 3.000000e-06*/
Related command	:TRIGger:DURATion:TLOWer

## :TRIGger:DURATion:TLOWer
Syntax	:TRIGger:DURATion:TLOWer <NR3> / :TRIGger:DURATion:TLOWer?
Description	Set or query the duration time lower limit in duration trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Real	8ns to 9.99s	1μs
Explanation	This command is available when the trigger condition (:TRIGger:DURATion:WHEN) is GREater or GLESs.
Return format	The query returns the duration time lower limit in scientific notation.
Example	:TRIGger:DURATion:TLOWer 0.000003 /*Set the duration time lower limit to 3μs*/ / :TRIGger:DURATion:TLOWer? /*The query returns 3.000000e-06*/
Related command	:TRIGger:DURATion:TUPPer

# :TRIGger:TIMeout (Option)

Command List:
- :TRIGger:TIMeout:SOURce
- :TRIGger:TIMeout:SLOPe
- :TRIGger:TIMeout:TIMe
## :TRIGger:TIMeout:SOURce
Syntax	:TRIGger:TIMeout:SOURce <source> / :TRIGger:TIMeout:SOURce?
Description	Set or query the trigger source in timeout trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:TIMeout:SOURce CHANnel2 /*Set the trigger source to CH2*/ / :TRIGger:TIMeout:SOURce? /*The query returns CHAN2*/

## :TRIGger:TIMeout:SLOPe
Syntax	:TRIGger:TIMeout:SLOPe <slope> / :TRIGger:TIMeout:SLOPe?
Description	Set or query the edge type in timeout trigger.
Parameter	Name	Type	Range	Default
	<slope>	Discrete	{POSitive|NEGative|RFALl}	POSitive
Explanation	POSitive: start timing when the rising edge of the input signal passes through the trigger level. / NEGative: start timing when the falling edge of the input signal passes through the trigger level. / RFALl: start timing when any edge of the input signal passes through the trigger level.
Return format	The query returns POS, NEG or RFAL.
Example	:TRIGger:TIMeout:SLOPe NEGative /*Set the edge type to falling edge*/ / :TRIGger:TIMeout:SLOPe? /*The query returns NEG*/
Related command

## :TRIGger:TIMeout:TIMe
Syntax	:TRIGger:TIMeout:TIMe <NR3> / :TRIGger:TIMeout:TIMe?
Description	Set or query the timeout time in timeout trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Real	16ns to 10s	16ns
Return format	The query returns the timeout time in scientific notation.
Example	:TRIGger:TIMeout:TIMe 0.002 /*Set the timeout time to 2ms*/ / :TRIGger:TIMeout:TIMe? /*The query returns 2.000000e-03*/

# :TRIGger:RUNT (Option)

Command List:
- :TRIGger:RUNT:SOURce
- :TRIGger:RUNT:POLarity
- :TRIGger:RUNT:WHEN
- :TRIGger:RUNT:WUPPer
- :TRIGger:RUNT:WLOWer
- :TRIGger:RUNT:ALEVel
- :TRIGger:RUNT:BLEVel
## :TRIGger:RUNT:SOURce
Syntax	:TRIGger:RUNT:SOURce <source> / :TRIGger:RUNT:SOURce?
Description	Set or query the trigger source in runt trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Return format	The query returns CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:RUNT:SOURce CHANnel2 /*Set the trigger source to CH2*/ / :TRIGger:RUNT:SOURce? /*The query returns CHAN2*/

## :TRIGger:RUNT:POLarity
Syntax	:TRIGger:RUNT:POLarity <polarity> / :TRIGger:RUNT:POLarity?
Description	Set or query the pulse polarity in runt trigger.
Parameter	Name	Type	Range	Default
	<polarity>	Discrete	{POSitive|NEGative}	POSitive
Explanation	POSitive: the instrument triggers on the positive runt pulse. / NEGative: the instrument triggers on the negative runt pulse.
Return format	The query returns POS or NEG.
Example	:TRIGger:RUNT:POLarity NEGative /*Set the pulse polarity to negative*/ / :TRIGger:RUNT:POLarity? /*The query returns NEG*/

## :TRIGger:RUNT:WHEN
Syntax	:TRIGger:RUNT:WHEN <when> / :TRIGger:RUNT:WHEN?
Description	Set or query the qualifier in runt trigger.
Parameter	Name	Type	Range	Default
	<when>	Discrete	{NONE|GREater|LESS|GLESs}	NONE
Explanation	NONE: do not set the trigger condition in runt trigger. / GREater: trigger when the runt pulse width is greater than the lower limit of the pulse width (refer to :TRIGger:RUNT:WLOWer). / LESS: trigger when the runt pulse width is lower than the upper limit of the pulse width (refer to :TRIGger:RUNT:WUPPer). / GLESs: trigger when the runt pulse width is greater than the lower limit (refer to :TRIGger:RUNT:WLOWer) and lower than the upper limit of the pulse width (refer to :TRIGger:RUNT:WUPPer). Note: the lower limit of the pulse width must be lower than the upper limit.
Return format	The query returns NONE, GRE, LESS or GLES.
Example	:TRIGger:RUNT:WHEN LESS /*Set the qualifier to <*/ / :TRIGger:RUNT:WHEN? /*The query returns LESS*/

## :TRIGger:RUNT:WUPPer
Syntax	:TRIGger:RUNT:WUPPer <NR3> / :TRIGger:RUNT:WUPPer?
Description	Set or query the pulse width upper limit in runt trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Discrete	16ns to 10s	2μs
Explanation	This command is only available when the qualifier (refer to :TRIGger:RUNT:WHEN) is LESS or GLESs.
Return format	The query returns the pulse width upper limit in scientific notation.
Example	:TRIGger:RUNT:WUPPer 0.02 /*Set the pulse width upper limit to 20ms*/ / :TRIGger:RUNT:WUPPer? /*The query returns 2.000000e-02*/
Related command	:TRIGger:RUNT:WLOWer

## :TRIGger:RUNT:WLOWer
Syntax	:TRIGger:RUNT:WLOWer <NR3> / :TRIGger:RUNT:WLOWer?
Description	Set or query the pulse width lower limit in runt trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Discrete	8ns to 9.99s	1μs
Explanation	This command is only available when the qualifier (refer to :TRIGger:RUNT:WHEN) is GREater or GLESs.
Return format	The query returns the pulse width lower limit in scientific notation.
Example	:TRIGger:RUNT:WLOWer 0.01 /*Set the pulse width lower limit to 1ms*/ / :TRIGger:RUNT:WLOWer? /*The query returns 1.000000e-03*/
Related command	:TRIGger:RUNT:WUPPer

## :TRIGger:RUNT:ALEVel
Syntax	:TRIGger:RUNT:ALEVel <level> / :TRIGger:RUNT:ALEVel?
Description	Set or query the trigger level upper limit in runt trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - OFFSet	2V
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet.
Return format	The query returns the trigger level upper limit in scientific notation.
Example	:TRIGger:RUNT:ALEVel 0.16 /*Set the trigger level upper limit to 160mV*/ / :TRIGger:RUNT:ALEVel? /*The query returns 1.600000e-01*/

## :TRIGger:RUNT:BLEVel
Syntax	:TRIGger:RUNT:BLEVel <level> / :TRIGger:RUNT:BLEVel?
Description	Set or query the trigger level lower limit in runt trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet.
Return format	The query returns the trigger level lower limit in scientific notation.
Example	:TRIGger:RUNT:BLEVel 0.16 /*Set the trigger level lower limit to 160mV*/ / :TRIGger:RUNT:BLEVel? /*The query returns 1.600000e-01*/

# :TRIGger:WINDows (Option)

Command List:
- :TRIGger:WINDows:SOURce
- :TRIGger:WINDows:SLOPe
- :TRIGger:WINDows:POSition
- :TRIGger:WINDows:TIMe
- :TRIGger:WINDows:ALEVel
- :TRIGger:WINDows:BLEVel
## :TRIGger:WINDows:SOURce
Syntax	:TRIGger:WINDows:SOURce <source> / :TRIGger:WINDows:SOURce?
Description	Set or query the trigger source in windows trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Return format	The query returns CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:WINDows:SOURce CHANnel2 /*Set the trigger source to CH2*/ / :TRIGger:WINDows:SOURce? /*The query returns CHAN2*/

## :TRIGger:WINDows:SLOPe
Syntax	:TRIGger:WINDows:SLOPe <type> / :TRIGger:WINDows:SLOPe?
Description	Set or query the windows type in windows trigger.
Parameter	Name	Type	Range	Default
	<type>	Discrete	{POSitive|NEGative|RFALl}	POSitive
Explanation	POSitive: trigger on the rising edge of the input signal when the voltage level is greater than the preset high trigger level. / NEGative: trigger on the falling edge of the input signal when the voltage level is lower than the preset low trigger level. / RFALl: trigger on any edge of the input signal when the voltage level meets the preset trigger level.
Return format	The query returns POS, NEG or RFAL.
Example	:TRIGger:WINDows:SLOPe NEGative /*Set the windows type to NEGative*/ / :TRIGger:WINDows:SLOPe? /*The query returns NEG*/

## :TRIGger:WINDows:POSition
Syntax	:TRIGger:WINDows:POSition <pos> / :TRIGger:WINDows:POSition?
Description	Set or query the trigger position in windows trigger.
Parameter	Name	Type	Range	Default
	<type>	Discrete	{EXIT|ENTER|TIMe}	ENTER
Explanation	EXIT: trigger when the input signal exits the specified trigger level range. / ENTER: trigger when the trigger signal enters the specified trigger level range. / TIMe: used to specify the hold time of the input signal after it enters the specified trigger level range. The instrument triggers when the accumulated hold time equals the windows time.
Return format	The query returns EXIT, ENTER or TIM.
Example	:TRIGger:WINDows:POSition ENTER /*Set the trigger position to enter*/ / :TRIGger:WINDows:POSition? /*The query returns ENTER*/

## :TRIGger:WINDows:TIMe
Syntax	:TRIGger:WINDows:TIMe <NR3> / :TRIGger:WINDows:TIMe?
Description	Set or query the hold time in windows trigger.
Parameter	Name	Type	Range	Default
	<NR3>	Real	8ns to 10s	1μs
Return format	The query returns the hold time in scientific notation.
Example	:TRIGger:WINDows:TIMe 0.002 /*Set the hold time to 2ms*/ / :TRIGger:WINDows:TIMe? /*The query returns 2.000000e-03*/
Related command	:TRIGger:WINDows:POSition

## :TRIGger:WINDows:ALEVel
Syntax	:TRIGger:WINDows:ALEVel <level> / :TRIGger:WINDows:ALEVel?
Description	Set or query the trigger level upper limit in windows trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - OFFSet	2V
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet.
Return format	The query returns the trigger level upper limit in scientific notation.
Example	:TRIGger:WINDows:ALEVel 0.16 /*Set the trigger level upper limit to 160mV*/ / :TRIGger:WINDows:ALEVel? /*The query returns 1.600000e-01*/

## :TRIGger:WINDows:BLEVel
Syntax	:TRIGger:WINDows:BLEVel <level> / :TRIGger:WINDows:BLEVel?
Description	Set or query the tigger level lower limit in windows trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet.
Return format	The query returns the tigger level lower limit in scientific notation.
Example	:TRIGger:WINDows:BLEVel 0.05 /*Set the tigger level lower limit to 50mV*/ / :TRIGger:WINDows:BLEVel? /*The query returns 5.000000e-02*/

# :TRIGger:DELay (Option)

Command List:
- :TRIGger:DELay:SA
- :TRIGger:DELay:SLOPA
- :TRIGger:DELay:SB
- :TRIGger:DELay:SLOPB
- :TRIGger:DELay:TYPe
- :TRIGger:DELay:TUPPer
- :TRIGger:DELay:TLOWer
## :TRIGger:DELay:SA
Syntax	:TRIGger:DELay:SA <source> / :TRIGger:DELay:SA?
Description	Set or query the trigger source A in delay trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Return format	The query returns CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:DELay:SA CHANnel2 /*Set the trigger source A to CH2*/ / :TRIGger:DELay:SA? /*The query returns CHAN2*/

## :TRIGger:DELay:SLOPA
Syntax	:TRIGger:DELay:SLOPA <slope> / :TRIGger:DELay:SLOPA?
Description	Set or query the edge type of edge A in delay trigger.
Parameter	Name	Type	Range	Default
	<slope>	Discrete	{POSitive|NEGative}	POSitive
Return format	The query returns POS or NEG.
Example	:TRIGger:DELay:SLOPA NEGative /*Set the edge type of edge A to falling edge*/ / :TRIGger:DELay:SLOPA? /*The query returns NEG*/

## :TRIGger:DELay:SB
Syntax	:TRIGger:DELay:SB <source> / :TRIGger:DELay:SB?
Description	Set or query the trigger source B in delay trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel2
Return format	The query returns CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:DELay:SB CHANnel4 /*Set the trigger source B to CH4*/ / :TRIGger:DELay:SB? /*The query returns CHAN4*/

## :TRIGger:DELay:SLOPB
Syntax	:TRIGger:DELay:SLOPB <slope> / :TRIGger:DELay:SLOPB?
Description	Set or query the edge type of edge B in delay trigger.
Parameter	Name	Type	Range	Default
	<slope>	Discrete	{POSitive|NEGative}	POSitive
Return format	The query returns POS or NEG.
Example	:TRIGger:DELay:SLOPB NEGative /*Set the edge type of edge B to falling edge*/ / :TRIGger:DELay:SLOPB? /*The query returns NEG*/

## :TRIGger:DELay:TYPe
Syntax	:TRIGger:DELay:TYPe <type> / :TRIGger:DELay:TYPe?
Description	Set or query the delay type in delay trigger.
Parameter	Name	Type	Range	Default
	<type>	Discrete	{GREater|LESS|GLESs |GOUT}	GREater
Explanation	GREater: trigger when the time difference (△T) between the specified edges of tigger source A and tigger source B is greater than the preset time limit. / LESS: trigger when the time difference (△T) between the specified edges of tigger source A and tigger source B is lower than the preset time limit. / GLESs: trigger when the time difference (△T) between the specified edges of tigger source A and tigger source B is greater than the lower limit of the preset time and lower than the upper limit of the preset time. Note that the time lower limit must be lower than the time upper limit. / GOUT: trigger when the time difference (△T) between the specified edges of tigger source A and tigger source B is lower than the lower limit of the preset time or greater than the upper limit of the preset time. Note that the time lower limit must be lower than the time upper limit.
Return format	The query returns GOUT, GRE, LESS or GLES.
Example	:TRIGger:DELay:TYPe GOUT /*Set the delay type to ><*/ / :TRIGger:DELay:TYPe? /*The query returns GOUT*/
Related commands	:TRIGger:DELay:TUPPer / :TRIGger:DELay:TLOWer

## :TRIGger:DELay:TUPPer
Syntax	:TRIGger:DELay:TUPPer <NR3> / :TRIGger:DELay:TUPPer?
Description	Set or query the upper limit of the delay time in delay trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Real	16ns to 10s	2μs
Explanation	This command is only available when the delay type (refer to :TRIGger:DELay:TYPe) is LESS, GOUT or GLESs.
Return format	The query returns the upper limit of the delay time in scientific notation.
Example	:TRIGger:DELay:TUPPer 0.002 /*Set the upper limit of the delay time to 2ms*/ / :TRIGger:DELay:TUPPer? /*The query returns 2.000000e-03*/
Related command	:TRIGger:DELay:TLOWer

## :TRIGger:DELay:TLOWer
Syntax	:TRIGger:DELay:TLOWer <NR3> / :TRIGger:DELay:TLOWer?
Description	Set or query the lower limit of the delay time in delay trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Real	Related to the delay type: GREater: 8ns to 10s / GOUT or GLESs: 8ns to 9.99s	1μs
Explanation	This command is only available when the delay type (refer to :TRIGger:DELay:TYPe) is GREater, GOUT or GLESs.
Return format	The query returns the lower limit of the delay time in scientific notation.
Example	:TRIGger:DELay:TLOWer 0.002 /*Set the lower limit of the delay time to 2ms*/ / :TRIGger:DELay:TLOWer? /*The query returns 2.000000e-03*/
Related command	:TRIGger:DELay:TUPPer

# :TRIGger:SHOLd (Option)

Command List:
- :TRIGger:SHOLd:DSrc
- :TRIGger:SHOLd:CSrc
- :TRIGger:SHOLd:SLOPe
- :TRIGger:SHOLd:PATTern
- :TRIGger:SHOLd:TYPe
- :TRIGger:SHOLd:STIMe
- :TRIGger:SHOLd:HTIMe
## :TRIGger:SHOLd:DSrc
Syntax	:TRIGger:SHOLd:DSrc <source> / :TRIGger:SHOLd:DSrc?
Description	Set or query the data source in setup/hold trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Return format	The query returns CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:SHOLd:DSrc CHANnel1 /*Set the data source to CH1*/ / :TRIGger:SHOLd:DSrc? /*The query returns CHAN1*/

## :TRIGger:SHOLd:CSrc
Syntax	:TRIGger:SHOLd:CSrc <source> / :TRIGger:SHOLd:CSrc?
Description	Set or query the clock source in setup/hold trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel2
Return format	The query returns CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:SHOLd:CSrc CHANnel2 /*Set the clock source to CH2*/ / :TRIGger:SHOLd:CSrc? /*The query returns CHAN2*/

## :TRIGger:SHOLd:SLOPe
Syntax	:TRIGger:SHOLd:SLOPe <slope> / :TRIGger:SHOLd:SLOPe?
Description	Set or query the edge type in setup/hold trigger.
Parameter	Name	Type	Range	Default
	<slope>	Discrete	{POSitive|NEGative}	POSitive
Return format	The query returns POS or NEG.
Example	:TRIGger:SHOLd:SLOPe NEGative /*Set the edge type to falling edge*/ / :TRIGger:SHOLd:SLOPe? /*The query returns NEG*/

## :TRIGger:SHOLd:PATTern
Syntax	:TRIGger:SHOLd:PATTern <pattern> / :TRIGger:SHOLd:PATTern?
Description	Set or query the data type in setup/hold trigger.
Parameter	Name	Type	Range	Default
	<pattern>	Discrete	{H|L}	H
Return format	H: high level / L: low level
Example	The query returns H or L.
Syntax	:TRIGger:SHOLd:PATTern L /*Set the data type to L*/ / :TRIGger:SHOLd:PATTern? /*The query returns L*/

## :TRIGger:SHOLd:TYPe
Syntax	:TRIGger:SHOLd:TYPe <type> / :TRIGger:SHOLd:TYPe?
Description	Set or query the setup type in setup/hold trigger.
Parameter	Name	Type	Range	Default
	<type>	Discrete	{SETup|HOLd|SETHOLd}	SETup
Explanation	SETup: the oscilloscope triggers when the setup time is less than the setting value (:TRIGger:SHOLd:STIMe). / HOLd: the oscilloscope triggers when the hold time is less than the setting value (:TRIGger:SHOLd:HTIMe). / SETHOLd: the oscilloscope triggers when the setup time or hold time is less than the corresponding setting value (:TRIGger:SHOLd:STIMe or :TRIGger:SHOLd:HTIMe).
Return format	The query returns SET, HOL or SETHOL.
Example	:TRIGger:SHOLd:TYPe SETHOLd /*Set the setup type to setup&hold*/ / :TRIGger:SHOLd:TYPe? /*The query returns SETHOL*/

## :TRIGger:SHOLd:STIMe
Syntax	:TRIGger:SHOLd:STIMe <NR3> / :TRIGger:SHOLd:STIMe?
Description	Set or query the setup time in setup/hold trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Real	8ns to 1s	1μs
Explanation	Setup time refers to the period of time that the data stays stable and constant before the clock signal of the trigger appears. / This command is only available when the setup type (refer to :TRIGger:SHOLd:TYPe) is SETup or SETHOLd.
Return format	The query returns the setup time in scientific notation.
Example	:TRIGger:SHOLd:STIMe 0.002 /*Set the setup time to 2ms*/ / :TRIGger:SHOLd:STIMe? /*The query returns 2.000000e-03*/
Related command	:TRIGger:SHOLd:HTIMe

## :TRIGger:SHOLd:HTIMe
Syntax	:TRIGger:SHOLd:HTIMe <NR3> / :TRIGger:SHOLd:HTIMe?
Description	Set or query the hold time in setup/hold trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Real	8ns to 1s	1μs
Explanation	Hold time refers to the period of time that the data stays stable and constant after the clock signal of the trigger appears. / This command is only available when the setup type (refer to :TRIGger:SHOLd:TYPe) is HOLd or SETHOLd.
Return format	The query returns the hold time in scientific notation.
Example	:TRIGger:SHOLd:HTIMe 0.002 /*Set the hold time to 2ms*/ / :TRIGger:SHOLd:HTIMe? /*The query returns 2.000000e-03*/
Related command	:TRIGger:SHOLd:STIMe

# :TRIGger:NEDGe (Option)

Command List:
- :TRIGger:NEDGe:SOURce
- :TRIGger:NEDGe:SLOPe
- :TRIGger:NEDGe:IDLE
- :TRIGger:NEDGe:EDGE
- :TRIGger:NEDGe:LEVel
## :TRIGger:NEDGe:SOURce
Syntax	:TRIGger:NEDGe:SOURce <source> / :TRIGger:NEDGe:SOURce?
Description	Set or query the trigger source in Nth edge trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| / D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:NEDGe:SOURce CHANnel2 /*Set the trigger source to CH2*/ / :TRIGger:NEDGe:SOURce? /*The query returns CHAN2*/

## :TRIGger:NEDGe:SLOPe
Syntax	:TRIGger:NEDGe:SLOPe <slope> / :TRIGger:NEDGe:SLOPe?
Description	Set or query the edge type in Nth edge trigger.
Parameter	Name	Type	Range	Default
	<slope>	Discrete	{POSitive|NEGative}	POSitive
Explanation	POSitive: trigger on the rising edge of the input signal when the voltage level meets the specified trigger level. / NEGative: trigger on the falling edge of the input signal when the voltage level meets the specified trigger level.
Return format	The query returns POS or NEG.
Example	:TRIGger:NEDGe:SLOPe NEGative /*Set the edge type to falling edge*/ / :TRIGger:NEDGe:SLOPe? /*The query returns NEG*/

## :TRIGger:NEDGe:IDLE
Syntax	:TRIGger:NEDGe:IDLE <NR3> / :TRIGger:NEDGe:IDLE?
Description	Set or query the idle time in Nth edge trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<NR3>	Real	16ns to 10s	1μs
Return format	The query returns the idle time in scientific notation.
Example	:TRIGger:NEDGe:IDLE 0.002 /*Set the idle time to 2ms*/ / :TRIGger:NEDGe:IDLE? /*The query returns 2.000000e-03*/

## :TRIGger:NEDGe:EDGE
Syntax	:TRIGger:NEDGe:EDGE <NR1> / :TRIGger:NEDGe:EDGE?
Description	Set or query the number of edges in Nth edge trigger.
Parameter	Name	Type	Range	Default
	<NR1>	Integer	1 to 65535	2
Return format	The query returns an integer between 1 and 65535.
Example	:TRIGger:NEDGe:EDGE 20 /*Set the number of edges to 20*/ / :TRIGger:NEDGe:EDGE? /*The query returns 20*/

## :TRIGger:NEDGe:LEVel
Syntax	:TRIGger:NEDGe:LEVel <level> / :TRIGger:NEDGe:LEVel?
Description	Set or query the trigger level in Nth edge trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - / OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet. / This command is only available when the signal source is an analog channel.
Return format	The query returns the trigger level in scientific notation.
Example	:TRIGger:NEDGe:LEVel 0.16 /*Set the trigger level to 160mV*/ / :TRIGger:NEDGe:LEVel? /*The query returns 1.600000e-01*/

# :TRIGger:RS232 (Option)

Command List:
- :TRIGger:RS232:SOURce
- :TRIGger:RS232:WHEN
- :TRIGger:RS232:PARity
- :TRIGger:RS232:STOP
- :TRIGger:RS232:DATA
- :TRIGger:RS232:WIDTh
- :TRIGger:RS232:BAUD
- :TRIGger:RS232:BUSer
- :TRIGger:RS232:LEVel
## :TRIGger:RS232:SOURce
Syntax	:TRIGger:RS232:SOURce <source> / :TRIGger:RS232:SOURce?
Description	Set or query the trigger source in RS232 trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| / D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:RS232:SOURce CHANnel2 /*Set the trigger source to CH2*/ / :TRIGger:RS232:SOURce? /*The query returns CHAN2*/

## :TRIGger:RS232:WHEN
Syntax	:TRIGger:RS232:WHEN <when> / :TRIGger:RS232:WHEN?
Description	Set or query the trigger condition in RS232 trigger.
Parameter	Name	Type	Range	Default
	<when>	Discrete	{STARt|ERRor|PARity|DATA}	STARt
Explanation	STARt: trigger on the start frame position. / ERRor: trigger when error frame is detected. / PARity: trigger when check error is detected. / DATA: trigger on the last bit of the preset data bits and even-odd check bits.
Return format	The query returns STAR, ERR, PAR or DATA.
Example	:TRIGger:RS232:WHEN ERRor /*Set the trigger condition to error frame*/ / :TRIGger:RS232:WHEN? /*The query returns ERR*/
Related command	:TRIGger:RS232:DATA

## :TRIGger:RS232:PARity
Syntax	:TRIGger:RS232:PARity <parity> / :TRIGger:RS232:PARity?
Description	Set or query the parity type when the trigger condition is ERRor or PARity in RS232 trigger.
Parameter	Name	Type	Range	Default
	<parity>	Discrete	{EVEN|ODD|NONE}	Error Frame: NONE / Check Error: ODD
Explanation	When the trigger condition is PARity, the parity type cannot be set to NONE. At this point, the default parity type is ODD.
Return format	The query returns EVEN, ODD or NONE.
Example	:TRIGger:RS232:PARity EVEN /*Set the parity type to even*/ / :TRIGger:RS232:PARity? /*The query returns EVEN*/
Related command	:TRIGger:RS232:WHEN

## :TRIGger:RS232:STOP
Syntax	:TRIGger:RS232:STOP <bit> / :TRIGger:RS232:STOP?
Description	Set or query the stop bit when the trigger condition is ERRor in RS232 trigger.
Parameter	Name	Type	Range	Default
	<bit>	Discrete	{1|2}	1
Return format	The query returns 1 or 2.
Example	:TRIGger:RS232:STOP 2 /*Set the stop bit to 2*/ / :TRIGger:RS232:STOP? /*The query returns 2*/
Related command	:TRIGger:RS232:WHEN

## :TRIGger:RS232:DATA
Syntax	:TRIGger:RS232:DATA <data> / :TRIGger:RS232:DATA?
Description	Set or query the data when the trigger condition is DATA in RS232 trigger.
Parameter	Name	Type	Range	Default
	<data>	Integer	0 to 2n-1	90
Explanation	In 2n-1, n is the current data bits and can be 5, 6, 7 or 8.
Return format	The query returns an integer.
Example	:TRIGger:RS232:DATA 10 /*Set the data to 10*/ / :TRIGger:RS232:DATA? /*The query returns 10*/
Related command	:TRIGger:RS232:WIDTh / :TRIGger:RS232:WHEN

## :TRIGger:RS232:WIDTh
Syntax	:TRIGger:RS232:WIDTh <width> / :TRIGger:RS232:WIDTh?
Description	Set or query the data bits when the trigger condition is DATA in RS232 trigger.
Parameter	Name	Type	Range	Default
	<width>	Discrete	{5|6|7|8}	8
Return format	The query returns 5, 6, 7 or 8.
Example	:TRIGger:RS232:WIDTh 6 /*Set the data bits to 6*/ / :TRIGger:RS232:WIDTh? /*The query returns 6*/
Related command	:TRIGger:RS232:WHEN

## :TRIGger:RS232:BAUD
Syntax	:TRIGger:RS232:BAUD <baud_rate> / :TRIGger:RS232:BAUD?
Description	Set or query the baud rate in RS232 trigger. The default unit is bps.
Parameter	Name	Type	Range	Default
	<baud_rate>	Discrete	{2400|4800|9600|19200|38400|57600| / 115200|230400|460800|921600|1000000| USER}	9600
Return format	The query returns an integer.
Example	:TRIGger:RS232:BAUD 4800 /*Set the baud rate to 4800*/ / :TRIGger:RS232:BAUD? /*The query returns 4800*/
Related command	:TRIGger:RS232:BUSer

## :TRIGger:RS232:BUSer
Syntax	:TRIGger:RS232:BUSer <user baud> / :TRIGger:RS232:BUSer?
Description	Set or query the user-defined baud rate in RS232 trigger. The default unit is bps.
Parameter	Name	Type	Range	Default
	<user baud>	Integer	110 to 20000000	9600
Return format	The query returns an integer.
Example	:TRIGger:RS232:BUSer 50000 /*Set the user-defined baud rate to 50000*/ :TRIGger:RS232:BUSer? /*The query returns 50000*/
Related command	:TRIGger:RS232:BAUD

## :TRIGger:RS232:LEVel
Syntax	:TRIGger:RS232:LEVel <level> / :TRIGger:RS232:LEVel?
Description	Set or query the trigger level in RS232 trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - / OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet. / This command is only available when the signal source is an analog channel.
Return format	The query returns the trigger level in scientific notation.
Example	:TRIGger:RS232:LEVel 0.16 /*Set the trigger level to 160mV*/ / :TRIGger:RS232:LEVel? /*The query returns 1.600000e-01*/

# :TRIGger:IIC (Option)

Command List:
- :TRIGger:IIC:SCL
- :TRIGger:IIC:SDA
- :TRIGger:IIC:WHEN
- :TRIGger:IIC:AWIDth
- :TRIGger:IIC:ADDRess
- :TRIGger:IIC:DIRection
- :TRIGger:IIC:DATA
- :TRIGger:IIC:CLEVel
- :TRIGger:IIC:DLEVel
## :TRIGger:IIC:SCL
Syntax	:TRIGger:IIC:SCL <source> / :TRIGger:IIC:SCL?
Description	Set or query the channel source of SCL in I2C trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:IIC:SCL CHANnel2 /*Set the SCL source to CH2*/ / :TRIGger:IIC:SCL? /*The query returns CHAN2*/

## :TRIGger:IIC:SDA
Syntax	:TRIGger:IIC:SDA <source> / :TRIGger:IIC:SDA?
Description	Set or query the channel source of SDA in I2C trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel2
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:IIC:SDA CHANnel2 /*Set the SDA source to CH2*/ / :TRIGger:IIC:SDA? /*The query returns CHAN2*/

## :TRIGger:IIC:WHEN
Syntax	:TRIGger:IIC:WHEN <trig_type> / :TRIGger:IIC:WHEN?
Description	Set or query the trigger condition in I2C trigger.
Parameter	Name	Type	Range	Default
	<trig_type>	Discrete	{STARt|RESTart|STOP|NACKnowledge| / ADDRess|DATA|ADATa}	STARt
Explanation	STARt: trigger when the SDA data transitions from high to low while the SCL is high. / RESTart: trigger when another start condition occurs before a stop condition. / STOP: trigger when the SDA data transitions from low to high while the SCL is high. / NACKnowledge: trigger when the SDA data is high during any acknowledgement of the SCL clock position. / ADDRess: search for the specified address value and trigger on the read/write bit. / DATA: search for the specified data value on the data line (SDA) and trigger on the jump edge of the clock line (SCL) corresponding to the last bit of the data. / ADATa:search for the specified address value and data value at the same time and trigger when the “Address” and “Data” conditions are met at the same time.
Return format	The query returns STAR, STOP, NACK, REST, ADDR, DATA or ADAT.
Example	:TRIGger:IIC:WHEN RESTart /*Set the trigger condition to restart*/ / :TRIGger:IIC:WHEN? /*The query returns REST*/

## :TRIGger:IIC:AWIDth
Syntax	:TRIGger:IIC:AWIDth <bits> / :TRIGger:IIC:AWIDth?
Description	Set or query the address bits when the trigger condition is ADDRess or ADATa in I2C trigger.
Parameter	Name	Type	Range	Default
	<bits>	Discrete	{7|8|10}	7
Explanation	The query returns 7, 8 or 10.
Return format	:TRIGger:IIC:AWIDth 10 /*Set the address bits to 10*/ / :TRIGger:IIC:AWIDth? /*The query returns 10*/
Example	:TRIGger:IIC:DIRection / :TRIGger:IIC:ADDRess

## :TRIGger:IIC:ADDRess
Syntax	:TRIGger:IIC:ADDRess <adr> / :TRIGger:IIC:ADDRess?
Description	Set or query the address when the trigger condition is ADDRess or ADATa in I2C trigger.
Parameter	Name	Type	Range	Default
	<adr>	Integer	0 to 2n – 1: 0 to 127, 0 to 255 or 0 to 1023	1
Explanation	In 2n – 1, n is the current address bits.
Return format	The query returns an integer.
Example	:TRIGger:IIC:ADDRess 100 /*Set the address to 100*/ / :TRIGger:IIC:ADDRess? /*The query returns 100*/
Related command	:TRIGger:IIC:AWIDth

## :TRIGger:IIC:DIRection
Syntax	:TRIGger:IIC:DIRection <dir> / :TRIGger:IIC:DIRection?
Description	Set or query the data direction when the trigger condition is ADDRess or ADATa in I2C trigger.
Parameter	Name	Type	Range	Default
	<dir>	Discrete	{READ|WRITe|RWRite}	READ
Explanation	This command is invalid when the address bits is 8.
Return format	The query returns READ, WRIT or RWR.
Example	:TRIGger:IIC:DIRection RWRite /*Set the data direction to read/write*/ / :TRIGger:IIC:DIRection? /*The query returns RWR*/
Related command	:TRIGger:IIC:AWIDth

## :TRIGger:IIC:DATA
Syntax	:TRIGger:IIC:DATA <data> / :TRIGger:IIC:DATA?
Description	Set or query the data when the trigger condition is DATA or ADATa in I2C trigger.
Parameter	Name	Type	Range	Default
	<data>	Integer	0 to 240 - 1	82
Explanation	The range of <data> is related to the byte length. The maximum byte length is 5, namely 40 bit binary data. Thus, the range of <data> is from 0 to 240 – 1.
Return format	The query returns an integer.
Example	:TRIGger:IIC:DATA 64 /*Set the data to 64*/ / :TRIGger:IIC:DATA? /*The query returns 64*/
Related command	:TRIGger:IIC:WHEN

## :TRIGger:IIC:CLEVel
Syntax	:TRIGger:IIC:CLEVel <level> / :TRIGger:IIC:CLEVel?
Description	Set or query the trigger level of SCL in I2C trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet. / This command is only available when the signal source of the clock line is an analog channel.
Return format	The query returns the trigger level of SCL in scientific notation.
Example	:TRIGger:IIC:CLEVel 0.16 /*Set the trigger level of SCL to 160mV*/ / :TRIGger:IIC:CLEVel? /*The query returns 1.600000e-01*/

## :TRIGger:IIC:DLEVel
Syntax	:TRIGger:IIC:DLEVel <level> / :TRIGger:IIC:DLEVel?
Description	Set or query the trigger level of SDA in I2C trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - / OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet. / This command is only available when the signal source of the data line is an analog channel.
Return format	The query returns the trigger level of SDA in scientific notation.
Example	:TRIGger:IIC:DLEVel 0.16 /*Set the trigger level of SDA to 160mV*/ :TRIGger:IIC:DLEVel? /*The query returns 1.600000e-01*/

# :TRIGger:SPI (Option)

Command List:
- :TRIGger:SPI:SCL
- :TRIGger:SPI:SDA
- :TRIGger:SPI:WHEN
- :TRIGger:SPI:WIDTh
- :TRIGger:SPI:DATA
- :TRIGger:SPI:TIMeout
- :TRIGger:SPI:SLOPe
- :TRIGger:SPI:CLEVel
- :TRIGger:SPI:DLEVel
- :TRIGger:SPI:SLEVel
- :TRIGger:SPI:MODE
- :TRIGger:SPI:CS
## :TRIGger:SPI:SCL
Syntax	:TRIGger:SPI:SCL <source> / :TRIGger:SPI:SCL?
Description	Set or query the channel source of SCL in SPI trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| / D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel1
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:SPI:SCL CHANnel1 /*Set the channel source of SCL to CH1*/ / :TRIGger:SPI:SCL? /*The query returns CHAN1*/

## :TRIGger:SPI:SDA
Syntax	:TRIGger:SPI:SDA <source> / :TRIGger:SPI:SDA?
Description	Set or query the channel source of SDA in SPI trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| / D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel2
Explanation	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Return format	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Example	:TRIGger:SPI:SDA CHANnel2 /*Set the channel source of SDA to CH2*/ / :TRIGger:SPI:SDA? /*The query returns CHAN2*/

## :TRIGger:SPI:WHEN
Syntax	:TRIGger:SPI:WHEN <trig_type> / :TRIGger:SPI:WHEN?
Description	Set or query the trigger condition in SPI trigger.
Parameter	Name	Type	Range	Default
	<trig_type>	Discrete	{CS|TIMeout}	CS
Explanation	When the trigger condition is set to TIMeout, you can use the :TRIGger:SPI:TIMeout command to set the timeout value.
Return format	The query returns CS or TIM.
Example	:TRIGger:SPI:WHEN TIMeout /*Set the trigger condition to TIMeout*/ :TRIGger:SPI:WHEN? /*The query returns TIM*/

## :TRIGger:SPI:WIDTh
Syntax	:TRIGger:SPI:WIDTh <width> / :TRIGger:SPI:WIDTh?
Description	Set or query the data bits of the SDA channel in SPI trigger.
Parameter	Name	Type	Range	Default
	<width>	Integer	4 to 32	8
Return format	The query returns an integer.
Example	:TRIGger:SPI:WIDTh 10 /*Set the data bits of the SDA channel to 10*/ / :TRIGger:SPI:WIDTh? /*The query returns 10*/
Related command	:TRIGger:SPI:DATA

## :TRIGger:SPI:DATA
Syntax	:TRIGger:SPI:DATA <data> / :TRIGger:SPI:DATA?
Description	Set or query the data in SPI trigger.
Parameter	Name	Type	Range	Default
	<data>	Integer	0 to 232-1	82
Explanation	The range of <data> is related to the data bits. The maximum data bits is 32. Thus, the range of <data> is from 0 to 232-1.
Return format	The query returns an integer.
Example	:TRIGger:SPI:DATA 5 /*Set the data to 5*/ / :TRIGger:SPI:DATA? /*The query returns 5*/
Related command	:TRIGger:SPI:WIDTh

## :TRIGger:SPI:TIMeout
Syntax	:TRIGger:SPI:TIMeout <time_value> / :TRIGger:SPI:TIMeout?
Description	Set or query the timeout value when the trigger condition is TIMeout in SPI trigger. The default unit is s.
Parameter	Name	Type	Range	Default
	<time_value>	Real	100ns to 1s	1μs
Return format	The query returns the timeout value in scientific notation.
Example	:TRIGger:SPI:TIMeout 0.001 /*Set the timeout value to 1ms*/ / :TRIGger:SPI:TIMeout? /*The query returns 1.000000e-03*/
Related command	:TRIGger:SPI:WHEN

## :TRIGger:SPI:SLOPe
Syntax	:TRIGger:SPI:SLOPe <slope> / :TRIGger:SPI:SLOPe?
Description	Set or query the clock edge in SPI trigger.
Parameter	Name	Type	Range	Default
	<slope>	Discrete	{POSitive|NEGative}	POSitive
Explanation	POSitive: sample the SDA data on the rising edge of the clock. NEGative: sample the SDA data on the falling edge of the clock.
Return format	The query returns POS or NEG.
Example	:TRIGger:SPI:SLOPe POSitive /*Set the clock edge to POSitive*/ / :TRIGger:SPI:SLOPe? /*The query returns POS*/

## :TRIGger:SPI:CLEVel
Syntax	:TRIGger:SPI:CLEVel <level> / :TRIGger:SPI:CLEVel?
Description	Set or query the trigger level of the SCL channel in SPI trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet. / This command is only available when the signal source of the clock line is an analog channel.
Return format	The query returns the trigger level of the SCL channel in scientific notation.
Example	:TRIGger:SPI:CLEVel 0.16 /*Set the trigger level of the SCL channel to 160mV*/ / :TRIGger:SPI:CLEVel? /*The query returns 1.600000e-01*/

## :TRIGger:SPI:DLEVel
Syntax	:TRIGger:SPI:DLEVel <level> / :TRIGger:SPI:DLEVel?
Description	Set or query the trigger level of the SDA channel in SPI trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet. / This command is only available when the signal source of the data line is an analog channel.
Return format	The query returns the trigger level of the SDA channel in scientific notation.
Example	:TRIGger:SPI:DLEVel 0.16 /*Set the trigger level of the SDA channel to 160mV*/ / :TRIGger:SPI:DLEVel? /*The query returns 1.600000e-01*/

## :TRIGger:SPI:SLEVel
Syntax	:TRIGger:SPI:SLEVel <level> / :TRIGger:SPI:SLEVel?
Description	Set or query the trigger level of the CS channel in SPI trigger. The unit is the same as the current amplitude unit.
Parameter	Name	Type	Range	Default
	<level>	Real	(± 5 × VerticalScale from the screen center) - / OFFSet	0
Explanation	For VerticalScale, refer to :CHANnel<n>:SCALe. For OFFSet, refer to :CHANnel<n>:OFFSet. / This command is only available when the signal source of the CS line is an analog channel.
Return format	The query returns the trigger level of the CS channel in scientific notation.
Example	:TRIGger:SPI:SLEVel 0.16 /*Set the trigger level of the CS channel to 160mV*/ / :TRIGger:SPI:SLEVel? /*The query returns 1.600000e-01*/

## :TRIGger:SPI:MODE
Syntax	:TRIGger:SPI:MODE <mode> / :TRIGger:SPI:MODE?
Description	Set or query the CS mode when the trigger condition is CS in SPI trigger.
Parameter	Name	Type	Range	Default
	<mode>	Discrete	{HIGH|LOW}	LOW
Return format	The query returns HIGH or LOW.
Example	:TRIGger:SPI:MODE LOW /*Set the CS mode to low*/ / :TRIGger:SPI:MODE? /*The query returns LOW*/
Related command	:TRIGger:SPI:WHEN

## :TRIGger:SPI:CS
Syntax	:TRIGger:SPI:CS <source> / :TRIGger:SPI:CS?
Description	Set or query the data source of the CS signal in SPI trigger.
Parameter	Name	Type	Range	Default
	<source>	Discrete	{D0|D1|D2|D3|D4|D5|D6|D7|D8| / D9|D10|D11|D12|D13|D14|D15| / CHANnel1|CHANnel2|CHANnel3|CHANnel4}	CHANnel3
Return format	When one or more of channels D0 to D7 is turned on, CH4 cannot be selected; when one or more of D8 to D15 is turned on, CH3 cannot be selected.
Example	The query returns D0, D1, D2, D3, D4, D5, D6, D7, D8, D9, D10, D11, D12, D13, D14, D15, CHAN1, CHAN2, CHAN3 or CHAN4.
Syntax	:TRIGger:SPI:CS CHANnel2 /*Set the data source of the CS signal to CH2*/ / :TRIGger:SPI:CS? /*The query returns CHAN2*/