# Common command reference
## *CLS
Syntax	*CLS
Description	Clear Status Command. This command clears all status data structures in the Instrument: / Standard Event Status Register / OPERation Event Status Register / QUEStionable Event Status Register / Error/Event Queue / The corresponding condition and enable registers are unaffected. If *CLS immediately follows a program message terminator (<NL>), then the output queue and the MAV bit are also cleared.
Return	None
Related Commands	*ESR? / STATus:OPERation[:EVENt] / STATus:OPERation:INSTrument[:EVENt] / STATus:OPERation:INSTrument:ISUMmary[<n>][:EVENt] / STATus:QUEStionable[:EVENt] / STATus:QUEStionable:INSTrument[:EVENt] / STATus:QUEStionable:INSTrument:ISUMmary[<n>][:EVENt] / SYSTem:ERRor

## *ESE
Syntax	*ESE {<value>} / *ESE?
Description	Standard Event Status Enable Command. This command sets the Standard Event Status Enable register bits in the Instrument. Those settings determine which events of the Standard Event Status Event register (see *ESR?) are allowed to set the ESB (Event Summary Bit) of the Status Byte register. A 1 in the bit position enables the corresponding event. All of the enabled events of the Standard Event Status Event Register are logically ORed sets the Event Summary Bit (ESB) of the Status Byte Register. / A STATus:PRESet command does not clear the bits in the Status Byte register.
Parameters	Name	Type	Range	Default
	<value>	NR1	0 – 255 (A decimal value which corresponds to the binary-weighted sum of the bits in the register)	None
Return	The query reads the enable register and returns a decimal value which corresponds to the binary-weighted sum of all bits set in the register.
Usage example	To enable bit 2 (decimal value = 4), bit 3 (decimal value = 8), and bit 7 (decimal value = 128), the corresponding decimal value would be 140 (4 + 8 + 128): / *ESE 140 / Read value of the Standard Event Status Enable register: / *ESE? / 140
Related Commands	*CLS / *ESR / *RST / STATus:PRESet

## *ESR?
Syntax	*ESR?
Description	Standard Event Status Register Query. Reading the Standard Event Status Event register clears it.
Return	The instrument returns a decimal value which corresponds to the binary-weighted sum of all bits in the register.
Usage example	If ERRor (bit 2) is set: / ESR? / 4

## *IDN?
Syntax	*IDN?
Description	Identification query for the UNIQUE identification of the instrument. (see also IEEE 488.2 10.14).
Return	The following system parameters will be displayed: <vendor>, <model>, <serial number>, <firmware>.
Usage example	*IDN? / Envox,EEZ H24005 (Due),CS17001,v1.1

## *OPC
Syntax	*OPC / *OPC?
Description	Operation Complete Command. The command is mainly used for program synchronization. It causes the instrument to set the OPC bit (bit 0) of the Standard Event Status register when the instrument has completed all pending operations *OPC. Pending operations are complete when: / All commands sent before *OPC is received, including paralleled commands, have been completed. Most commands are sequential and are completed before the next command is executed. Commands that affect output voltage, current, or state, relays, and trigger actions are executed in parallel with subsequent commands. *OPC provides notification that all parallel commands have completed. / All triggered actions are completed. / Query whether the current operation is completed and the query returns 1. / See also IEEE 488.2 Section 12.5 – 12.8.
Return	Query causes the instrument to place a 1 in the output buffer when all pending operations are completed. *OPC? does not suspend processing of commands.
Usage example	*OPC? / 1 / if current operation is not completed: / *OPC? / 0

## *RCL
Syntax	*RCL {<profile>}
Description	This command recalls the instrument state stored in the specified storage location. / It is not possible to recall the instrument state from a storage location that is empty or was deleted.
Parameters	Name	Type	Range	Default
	<profile>	NR1	0 – 9	None
Return	None
Usage example	*RCL 2
Related Commands	*SAV

## *RST
Syntax	*RST
Description	Reset Command. Restores the instrument to its initial state (as predefined in the instrument firmware) and clears the error queue.
Return	None
Usage example	*RST
Related Commands	*RST / *SAV

## *SAV
Syntax	*SAV {<profile>}
Description	This command stores the current instrument state in the specified storage location. Any state previously stored in the same location is overwritten without generating any errors. / A reset (*RST command) does not affect the configurations stored in memory. Once a state is stored, it remains constant until it is overwritten using this command.
Parameters	Name	Type	Range	Default
	<profile>	NR1	1 – 9	None
Return	None
Usage example	*SAV 2
Related Commands	*RCL / *RST

## *SRE
Syntax	*SRE {<value>} / *SRE?
Description	Enable bits in the Status Byte enable register.
Parameters	Name	Type	Range	Default
	<value>	NR1	0 – 255	None
Return	Query the Status Byte enable register. The instrument returns a decimal value which corresponds to the binary-weighted sum of all bits set in the enable register.
Usage example	*SRE 32
Related Commands	*STB

## *STB?
Syntax	*STB?
Description	Read Status Byte Query. This query reads the Status Byte register, which contains the status summary bits and the Output Queue MAV bit. The Status Byte is a read-only register and its bits are not cleared when it is read. / A serial poll also returns the value of the Status Byte register, except that bit 6 returns Request for Service (RQS) instead of Master Status Summary (MSS). A serial poll clears RQS, but not MSS. When the MSS bit is set, the instrument has one or more reasons for requesting service.
Return	The instrument returns a decimal value which corresponds to the binary-weighted sum of all bits in the register.
Usage example	If OPER (bit 7) is set: / *STB? / 128
Related Commands	*SRE

## *TRG
Syntax	*TRG
Description	This command generates a trigger to the trigger subsystem which has selected a bus / (software) trigger as its source.
Return	None
Usage example	Generate a trigger operation after 5 seconds: / TRIG:SOUR BUS / TRIG:DEL 5 / INIT / *TRG
Related Commands	*WAI

## *TST?
Syntax	*TST?
Description	Self-Test Query. The self-test query causes an internal self-test, and places a response into the Output Queue indicating whether or not the instrument completed the self-test without detected errors.
Return	0 or 1 depends of the self-test results.
Usage example	If all tests passed: / *TST? / 0 / If one or more tests failed: / *TST? / 1

## *WAI
Syntax	*WAI
Description	The Wait-to-Continue Command causes the instrument to wait until all pending commands are completed before executing any other commands. / Pending operations are as defined under the *OPC command.
Return	None
Usage example	For example, the *WAI command can be used to make a voltage measurement after an OUTPut ON command has completed: / OUTPut ON;*WAI;:MEASure:VOLTage?
Related Commands	*OPC